[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne R23-MED (0110023)
O2 Sensor Test Procedure

VM3/COP/37.38
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



NB R-23Med & R-23V are identical sensors from different manufacturers

Test & Equipment Required

Functional test required using R15 / R23 tester and sensor holder, (2 x digital multimeters & adapter box).

Output specification as in document: (VM3COP37.00, column (i)). Test box converts current output to 17mV ± 3mV

Method.

1. Open the packaging and remove the O2 sensor.

2. Leaving shorting gauze and clip in place check the O2 sensor for damage and signs of leaking electrolyte.

3. Check for an output and wait for both meter readings to stabilize may take a minute

4. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 5 minutes - Maximum.

5. Remove shorting gauze and clip Ensure that both meter readings lay within the output specification range,  to 17mV ± 3mV. If the O2 sensor fails the output specification test, it should be resealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with SRN

6. Remove the O2 sensor from the sensor holder, refit shorting gauze and clip and label with Viamed stickers if appropriate.

7. Reseal in serial numbered packet and re-box. Integrity-seal the edge of the packet with initialed and dated Viamed sticker 
















VM3COP 37.38 Teledyne-R23 & V23.doc                  28/07/2006                 Page 1 of 1

[image: image1.png]