[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       
RESUSCITATION BED   – CE FILE 


Risk Assessment iaw EN ISO 14971:2000 Annex D : Possible hazards with medical devices.

Ref.
Hazard.


Related part / Component posing risk.
Sev

of

Haz.
Like

of

Haz.
Risk.
Solution.
Document referenced.
Sev

of

Haz.
Like

0f

Haz.
Risk.

D.2
Energy hazards and contributory factors

D.2.1
Electricity
N/A
1
1
1


1
1
1

D.2.2
Heat

1
1
1

(YZ) Design File
1
1
1

D.2.3
Mechanical force
N/A
2
1
2


2
1
2

D.2.4
Ionising radiation
N/A
1
1
1


1
1
1

D.2.5
Non ionising radiation
N/A
1
1
1


1
1
1

D.2.6
Moving parts
N/A
1
1
1


1
1
1

D.2.7
Unintended motion
N/A
1
1
1


1
1
1

D.2.8
Suspended masses
N/A
1
1
1


1
1
1

D.2.9
Patient support failure
N/A
1
1
1


1
1
1

D.2.10
Pressure (vessel rupture)
N/A
1
1
1


1
1
1

D.2.11
Acoustic pressure
N/A
1
1
1


1
1
1

D.2.12
Vibration
N/A
1
1
1


1
1
1

D.2.13
Magnetic fields (eg.MRI)
N/A
1
1
1

F. User Instructions
1
1
1

D.3
Biological hazards and contributory factors

D.3.1
Bio-contamination
Mattress
2
2
4
Easy to clean – cleaning instructions in user instructions
F. User Instructions
2
1
2

D.3.2
Bio-incompatibility
Mattress
1
1
1

Manufacturers Data
1
1
1

D.3.3
Incorrect formulation (chemical composition)
N/A
1
1
1


1
1
1

D.3.4
Toxicity
N/A
1
1
1

Manufacturers Data
1
1
1

D.3.5
Allergenicity
N/A
1
1
1

Manufacturers Data
1
1
1

D.3.6
Mutagenicity
N/A
1
1
1

Manufacturers Data
1
1
1

D.3.7
Oncogenicity
N/A
1
1
1

Manufacturers Data
1
1
1

D.3.8
Carcinogenicity
N/A
1
1
1

Manufacturers Data
1
1
1

D.3.9
Re and/or cross infection
Instrument & Leads
1
1
1

Manufacturers Data
1
1
1

D.3.10
Pyrogenicity
N/A
1
1
1

Manufacturers Data
1
1
1

D.3.11
Inability to maintain hygienic standards
Instrument & Leads
2
2
4
Easy to clean – cleaning instructions in user manual
F. User Instructions
2
1
2

D.3.12
Degradation
Instrument & Leads
2
1
2
Care Instructions given in user manual
F. User Instructions
1
1
1

D.4
Environmental hazards and contributory factors

D.4.1
Electromagnetic fields
N/A
1
1
1


1
1
1

D.4.2
Susceptibility to electromagnetic interference
N/A
1
1
1


1
1
1

D.4.3
Emissions of electromagnetic interference 
N/A
1
1
1


1
1
1

D.4.4
Inadequate supply of power 
N/A
1
1
1


1
1
1

D.4.5
Inadequate supply of coolant
N/A
1
1
1


1
1
1

D.4.6
Storage / operation outside prescribed environmental conditions
Must be locked
2
1
2
Label

2
1
2

D.4.7
Incompatibility with other devices with which the product is intended to be used

1
1
1


1
1
1

D.4.8
Accidental mechanical damage
Bed not locked when the door is closed
2
1
2
Relatively robust material used. If damaged, user to assess level of damage – spares parts available
F. User Instructions
1
1
1

D.4.9
Contamination due to waste products and/or device disposal
s
1
1
1


1
1
1

D.5
Hazards resulting from incorrect output of energy and substances

D.5.1
Electricity
N/A
1
1
1


1
1
1

D.5.2
Radiation
N/A
1
1
1


1
1
1

D.5.3
Volume
N/A
1
1
1


1
1
1

D.5.4
Pressure
N/A
1
1
1


1
1
1

D.5.5
Supply of medical gases
N/A
1
1
1


1
1
1

D.5.6
Supply of anaesthetic agents
N/A
1
1
1


1
1
1

D.6
Hazards related to the use of the medical device and contributory factors

D.6.1
Inadequate labelling
Label
2
1
2
Product easy to use.

Label
F. Labels
1
1
1

D.6.2
Inadequate operating instructions
User Instructions
2
1
2
Product easy to use.

User manual
F. User Instructions
1
1
1

D.6.3
Inadequate specification of accessories
N/A
1
1
1


1
1
1

D.6.4
Inadequate specification of pre-use checks
User Instructions / Labels
2
2
4
Product easy to use.

User manual / labels
F. User Instructions

F. Labels
2
1
2

D.6.5
Over-complicated operating instructions
User Instructions
2
1
2
Product easy to use.

User manual
F. User Instructions
1
1
1

D.6.6
Inadequate specification of service and maintenance
N/A
1
1
1
Normal PPM required

1
1
1

D.6.7
Use by unskilled / untrained personnel

3
2
6
Product easy to use. 
F. User Instructions
2
1
2

D.6.8
Reasonable foreseeable misuse

3
2
6

F. User Instructions
2
1
2

D.6.9
Insufficient warning of side effects
N/A
1
1
1


1
1
1

D.6.10
Inadequate warnings of hazards likely with re-use of single use devices
N/A
1
1
1


1
1
1

D.6.11
Incorrect measurement and other metrological aspects
N/A
1
1
1


1
1
1

D.6.12
Misrepresentation of results

1
1
1


1
1
1

D.6.13
Incompatibility with consumables / accessories / other devices

1
1
1


1
1
1

D.6.14
Sharp edges or points

2
2
4
Only if damaged, user to assess level of damage – spares parts available
F. User Instructions
1
1
1

D.7
Inappropriate, inadequate or overcomplicated user interface (man/machine communication)

D.7.1
Mistakes & judgement errors
N/A
1
1
1
No complex user interface

1
1
1

D.7.2
Lapses and cognitive recall errors
N/A
1
1
1
No complex user interface

1
1
1

D.7.3
Slips & blunders (mental or physical)
N/A
1
1
1
No complex user interface

1
1
1

D.7.4
Violation or abbreviation of instructions,  procedures etc
N/A
1
1
1
No complex user interface

1
1
1

D.7.5
Complex or confusing control system
N/A
1
1
1
No complex user interface

1
1
1

D.7.6
Ambiguous or unclear device state
N/A
1
1
1
No complex user interface

1
1
1

D.7.7
Ambiguous or unclear presentation of settings, measurement, or other information
N/A
1
1
1
No complex user interface

1
1
1

D.7.8
Misrepresentation of results
N/A
1
1
1
No results presented

1
1
1

D.7.9
Insufficient visibility, audibility or tactility
N/A
1
1
1
No complex user interface

1
1
1

D.7.10
Poor mapping of controls to action or of displayed information to actual state
N/A
1
1
1
No complex user interface

1
1
1

D.7.11
Controversial modes or mappings as compared to existing equipment
N/A
1
1
1
No complex user interface

1
1
1

D.8
Hazards arising from function failure, maintenance and ageing and contributory factors

D.8.1
Erroneous data transfer
N/A
1
1
1


1
1
1

D.8.2
Lack of, or inadequate specification for maintenance including post maintenance functional tests
N/A
1
1
1


1
1
1

D.8.4
Inadequate maintenance
N/A
1
1
1


1
1
1

D.8.5
Lack of adequate determination of end of device life

1
1
1
User decision based on damage or cost effectiveness of repair

1
1
1

D.8.6
Loss of electrical integrity

1
1
1


1
1
1

D.8.7
Loss of mechanical integrity
1
1
1
1
User decision based on damage or cost effectiveness of repair

1
1
1

D.8.8
Inadequate packaging (contamination and / or deterioration of the device)
1
2
2
4
1
M. Packaging
1
1
1

D.8.9
Re-use and/or improper re-use
1
1
1
1
User decision on suitability for next use based on pre-use check & tolerable damage 

1
1
1

D.8.10
Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
1
1
1
1
1

1
1
1

E. Risk Management BS14971 Annex D_Microstim.doc
    Page 1



14/05/2008
Severity of hazard: 1 = Insignificant, 2 =  Tolerable, 3  = Critical, 4  =  Intolerable.  Probability of event: 1 = Improbable, 2 = Occasional, 3 = Likely, 4 = Highly likely.  

Risk calculated as severity of hazard x probability of event, 1 – 16.  Further explanation of risk management policy – see Risk Management Policy & Definitions.


