[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne R24-MED (0110024), R24SIE (0110324)

O2 Sensor Test Procedure

VM3/COP/37.39

	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



NB R24SIE is a re-badged R-24MED

Test & Equipment Required

Functional test required using R22/R24 tester and R22/R24 test lead, (digital multimeter & 3 pin molex adapter cable).

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Connect the O2 sensor to the R22/R24 tester, select 200mV range (marked with arrows) and wait for meter reading to stabilise. Check for an output. Ensure that meter reading lies within the output specification range,  10mV- 17mV 

4. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hours.


5. If the O2 sensor fails the output specification test, it should be left  unsealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN.

6.  Record the voltage on a record sheet so as to be able to match two together for pairing.

7. Disconnect the O2 sensor from the R22/R24 tester and label with Viamed stickers if appropriate.

8. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box. 
















VM3COP 37.39 Teledyne-R24MED & SIE.doc      28/07/2006       Page 1 of 1

[image: image1.png]