[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


R29IMED (0110035)
O2 Sensor Test Procedure

VM3/COP/37.41
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



Test & Equipment Required

Functional test required using R29 tester and (digital voltmeter).

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte. 

3. Connect the O2 sensor to the R29 tester, select 200mV range (marked with arrows) and wait for a few seconds for the meter reading to stabilise Check for an output If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hour..

4. Ensure that meter reading lies within the output specification range, 10mV – 15.5mV .
If the O2 sensor fails the output specification test, it should left unsealed and boxed, placed in a ducket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN.

5. Disconnect the O2 sensor from the R29 tester and label with Viamed stickers if appropriate.

6. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box.















VM3COP 37.41 Teledyne-R29IMED.doc      28/07/2006      Page 1 of 1

[image: image1.png]