[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne R33D (0110214), R33D1 (0110215), R33VAN
O2 Sensor Test Procedure

VM3/COP/37.45
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



Test & Equipment Required

Visual check  Digital voltmeter  R22 adapter.

Output specification as in document: (VM3COP37.00, column (i)). 25mV +/- 2mV 

Method.

1. Open the packaging and remove the O2 sensor


2. Check the O2 sensor damage and signs of leaking electrolyte.

3. Check for an output  Ensure that  reading lies within the output specification range,
25mV +/- 2mV

4. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hours Minimum 

5. If the O2 sensor fails the output specification test, it should left unsealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN..

6. Label with Viamed stickers if appropriate.

7. Reseal in serial numbered packet and re-box. Integrity-seal the edge of the packet with initialed and dated Viamed sticker.
















VM3COP 37.45 Teledyne-R33D  all .doc      28/07/2006                    Page 1 of 1

[image: image1.png]