[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne T2 (0110052)
O2 Sensor Test Procedure

VM3/COP/37.47
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



Test & Equipment Required

Functional test using T1/T2 tester (analogue meter).

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open cell packaging and remove O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Remove the shorting plug, fit the O2 sensor into the T1/T2 tester and wait for the meter needle to stabilise. Check for an output If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hours

4. Ensure that meter reading can be adjusted to read 21% on the scale using the calibration control knob. If the O2 sensor fails the output specification test, it should be resealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN.

5. Remove the O2 sensor from the T1/T2 tester and refit the shorting plug.

6. Label with Viamed stickers if appropriate.

7. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box. 
















VM3COP 37.47 Teledyne-T2.doc      28/07/2006      Page 1 of 1

[image: image1.png]