[image: image1.png]


 

                   O2 Analysers

	COMPANY OPERATING PROCEDURES


TED191 (PN 0111191)
O2 Analyser Test Procedure

VM3/COP/38.01
	Date:
  30-Jan-03
	Revision date: 30-Jan-03
	Issue:   3



Test Equipment Required 

100% O2 supply, flow tube & Tee piece.

Method

1. Take the TED191 and T7 O2 sensor out of their packaging and examine for signs of physical damage and leaking electrolyte from the sensor. 

2. Fit the batteries and switch the TED191 on.

3. Ensure that the display shows 00 +/- 1% oxygen concentration. (1)
4. Switch the unit off.

5. Connect the T7 O2 sensor to the TED191.  Install the T7 O2 sensor with flow diverter fitted into a 100% oxygen circuit using the Tee piece. Adjust the flow rate of O2 to 5 litres / minute and when a stable displayed O2 reading is obtained on the TED191, adjust the calibration control to obtain a displayed O2 reading of 100 ( 1%. (2)
6. Turn the 100% O2 off and remove the T7 O2 sensor and flow diverter.  Unscrew the flow diverter and ensure that the displayed O2 reading gradually reduces to 21 ( 1%. (2)
7. If necessary adjust the calibration control to obtain a displayed O2 reading of 21%.

8. Test the upper alarm & threshold – turn the “set upper alarm” control until the upper alarm threshold is below 21%.  Ensure that the upper alarm sounds (at two beeps/sec) and visual alarm indicator activates (at 2 flashes/sec).  Ensure that the alarm cancel/delay operates correctly by pressing the silence alarm button and waiting the specified time until the TED191 re-alarms. Return the “set upper alarm” control to the fully clockwise position. (1)
9. Test the lower alarm & threshold – turn the “set lower alarm” control until the lower alarm threshold is above 21%.  Ensure that the lower alarm sounds (at 1 beep/sec) and that the visual alarm indicator activates (at 1 flash/sec). Ensure that the alarm cancel/delay operates correctly by pressing the silence alarm button and waiting the specified time until the TED191 re-alarms.  Return the “set lower alarm” control to the fully anticlockwise position. (1) 

10. Switch off.

11. Disconnect the T7 O2 sensor and cable.  Repackage as per the T7 O2 sensor test procedure (VM3COP37.52).

12. Remove the batteries from the TED191.

13. Label the TED191 with Viamed stickers if necessary.

14. Repackage the TED191 and accessories.

15. Label the box as necessary for UK or export.

16. Ensure that the TED191 and T7 O2 sensor are booked out in the stock books.

(1) 
If the unit fails to show a suitable reading at this stage the TED191 is faulty.  The unit and accessories should be repackaged and left at goods-in, in a docket with a note attached.

(2) 
If the unit fails to show a suitable reading at this stage the TED191 or T7 sensor is faulty.  The unit and accessories should be repackaged and left at goods-in, in a docket with a note attached.















Page 2 of 2

[image: image1.png]