[image: image1.jpg]


VANDAGRAPH

DIN KIT

DIN KIT Assembly Procedure

VM3/COP/40.03

	Date: 23 Sep. 03
	Revision Date: 17-Jul-06
	Issue: 1



	Parts Required


	Quantity
	Description
	Part No.
	

	1
	A268 Tee Adapter (blue)
	9711001
	

	
	22mm (M), 22mm(F), 15mm(F)
	
	

	1
	DMM22M10 Male Tubing Adapter (clear)
	9711004
	

	
	22mm (M) to 10mm
	
	

	1
	B-50057 Diverter
	9711002
	

	1
	BS111 Viton “O” ring – Oxygen compatible
	9711003
	

	1
	DIN 22F High Pressure Restrictor
	9711006
	

	1
	VP12 Tubing – 30cm long
	9710040
	

	1
	Leaflet
	SL002/9909
	

	1
	Clear bag (6inch x 9inch)
	PBSSBW11
	

	1
	Label
	
	


[image: image2.jpg]



Each DIN KIT comprises of the above components in the Parts List. The complete Kit is placed in the clear bag with a “Leaflet” folded so that the Specification is visible, and the DIN KIT lies on top of the description.

	Assembly:


1.0 Fit the Flo-Divertor to the bottom of the A268 “T” adapter, see fig 1.1.

       Fig 1.1

[image: image3.jpg]0-100% oxy

=% of ful

periiiog emperatre Tange: | |

BaHTe e |
i Weight:

Cable Length:

Flow-Thru Diverto

Tubing 30cm

8 ( ambient temperature
fienaline drift. 1 the fIOW
SINKIT metr od

so that the
s 100 Weh
qs fast an




[image: image4.jpg]



2.0 Measure, and cut, the VP12 tubing, soak the end of the tubing in a few cm’s of warm water to soften it, then fit to the DM22M Tubing adapter (if the water is too hot, or the tube is soaked too long, the tubing will discolour.

3.0 Examine the DIN22 High Pressure Restrictor for the following:

[image: image5.jpg]



Fig 1.2

a) Shards of Chrome in the channel (if any is evident – remove gently with a dentists pick, tooth pick or other such sharp instrument).

b) Any cosmetic fault with the Chroming.

c) Check that there is a small hole visible through the neck.

4.0 Place the “O”-ring over the channel and press home using tool.

[image: image6.jpg]




Fig 1.3

5.0 Test the screw thread by attaching to an “A” Clamp, 

6.0 Fit the DIN22 Restrictor to the 22mm end of the “T” piece.

7.0 Fit the DM22M to the male end of the “T” piece with the curve of the tubing pointing in the direction of the leg of the “T”, see fig 1.4.


Fig 1.4

	Packaging


1.0 Fold the specification leaflet with the specifications showing.

2.0 Place the leaflet in the clear bag with the opening to the right of the script in the readable position, see fig 1.5.


Fig 1.5

3.0 Bend the tubing towards the “T” leg and insert the whole into the bag with the tubing at the top, see fig 1.5.




�





Equipment Required:








�





�





�





�





�








VM3COP40.03 DINKIT Assembly 060315         July 17, 2006      Page 1 of 3

