S4000 NIBP Simulator Calibration Record Sheet.

Unit Serial No.


Engineer.


Date.



1.  Regulated power supply checks.

Tp1.  (Expect +5.0  0.25V): 
Tp2.  (Expect +12.0  0.3V):


Tp3.  (Expect -12.0  0.3V): 


2.  Pressure sensor reference voltage checks.

Tp5.    (Expect 5  0.001V). 
1st check:   
After adjustment:

Tp6.    (Expect 0  0.001V).
1st check: 
After adjustment:

3.  Dynamic pneumatic transducer drive signal checks.

Tp8.  Pneumatic transducer drive signal, post amp. Expect:3.3  0.3VP-P at 1.33Hz / 80bpm.

1st check.


After adjustment.

Not required.



Tp9.  Pneumatic transducer drive signal, pre amp. Expect:  0.1VP-P at 1.33Hz / 80bpm.

1st check.


After adjustment.

Not required.



4.  ECG drive signal checks.

Tp7. ECG drive signal, post amp: Expect : 2.0  0.2VP-P at 2.66Hz / 80bpm.

1st check.


After adjustment.

Not required.

Tp10.  ECG drive signal, pre amp: Expect 4.8  0.3VP-P at 2.66Hz / 80bpm.

1st check.


After adjustment.

Not required.

5.  Pneumatic system integrity checks.

Pressure: 300 mmHg.
1st check: 
After adjustment:

Pressure: 200 mmHg.
1st check: 
After adjustment: 

Pressure: 100 mmHg.
1st check: 
After adjustment:

Leakage at 200mmHg.



6.  Oven temperature check.

Oven: 37.4 0.2°C.
1st check:
After adjustment:

7.  Miscellaneous.

