[image: image1.jpg]Sensor R-17 VAN
Batt: MN1500 x2

Ser. No. TX




VANDAGRAPH

TEK-OX 

TEK-OX Assembly Procedure

VM3/COP/42.01

Date: 11-Aug-2008
Revision Date: 11-Aug-08
Issue: 2


Quantity
Description
Part No.



1
TEK-OX Main Body Insert Black 
9530003


1
TEK-OX Black  Front Cover
9530005


1
TEK-OX Black  Bottom cover
9530004


1
TEK-OX PC Lens

9530009


1
TEK-OX Black Shell left
9530012


1
TEK-OX Black  Shell Right
9530013


1
Blu-Ox Base Seal
9530002


1
Blu-Ox Front gasket
9530006


1
Blu-Ox  switch
9730036


1
Blu-Ox Cal contro
9730040


1
Blu-Ox Photo thread
9530010


1
Blu-Ox key ring
9530008


1
Blu-Ox Inside label
9530007


1
Blu-Ox Instruction manual
9590001


1
Vandagraph Logo round
9530001


1
TEX-OX Front Label Black
9530002


1
Blu-Ox  switch waterproof kit
9730037


1
Blu-Ox battery terminals right
9730064


1
Blu-Ox battery terminals left
9730063


1
Blu-Ox battery terminal common
9730062


1
LCD assembly
3501002


1
PCB assembly
9530015


1
Jack plug - plastic 750 1/8"
9071007


1
5/16” UNF Nuts
9520001


1
Cable assembly
9070004


2
Blu-Ox batteries AA
9060005


1
R-17VAN



2
Screws Black  M2 x 5 mm
9730061


5
Screws Black  M2 x 10 mm
9730060


1
Double sided tape
9711017


`1
Locktite 
9073005


1
Silicon grease
9711018


1
Quick setting adhesive (UHU)
9711020


1
Slicone rubber black non corrosive CAF10 
9711000


1
Conformal coating RS 714-462



1
Blu-Tac or equivalent



[image: image2.jpg]TEK-OX Oxygen Analyser
VANDAGRAPH LTD. d
15 Station Read, Gross Hills

Keighlay, W.Yarkshirz, BD20 70T, UK
Tel. +44 (0)1535 624300

Fax. +44 {0)1535 6355582 c €
Sen Na. TX _—




Warning   Glue can effect the plastic. Use sparingly and remove any surplus immediately

   [image: image3.jpg]



Assembly:

1. Programme the PCB ( See Appendix A)[image: image4.jpg]



2. Locate the Battery cover on the plastic pillar with the pips on top  and ensure the cut-out faces the screw hole. Remove plastic pip leaving 2mm . Using hot iron melt the pip to seal in the contacts

3. Clean the terminals with Isopropyl Alcohol.

4. Lubricate the battery compartment seal and insert into the battery compartment cover

5. Tin the battery terminals 

6. Locate the battery springs correctly the springs are offset[image: image5.jpg]


 “Left to left”  & “Right to right”.

7. Glue into position applying glue to the both parts of the right angle bend of connector, inside face . Glue / silicone or another compound can be used. The criteria is to seal the battery compartment from the electronics without surplus compound covering the spring terminals  In the centre of the connector spring is a hole which locates on a pip in the case

8. The battery terminals inside the case should be sealed with black CAF 10.or equivalent .  NB This is non corrosive.

9. The switch should then  be fitted 1st

10. Add switch into the enclosure: mate the locating peg and add the bridge sealing kit,

11. This has to be tightened through the right hand enclosure aperture

12. Add O ring to the  potentiometer and feed pot into enclosure body .

13. The  White spot on the pot should be at the top when the pot is half turned.

14. Add Locking washer ,Permabond (Loctite) or glue and tighten nut.
NB If the nut is difficult to locate turn it over NB some nuts are problematical

15. Tin both contacts of the Jack plug. Try the nut to ensure it screws to the end.NB some nuts are problematical

16. Add glue to the Jack plug and insert into the body. Ensure the plug is in the centre and the large terminal ( screen) is preferably at the back. Fasten with the nut and use loctite. This nut should be siliconed to stop vibration breaking the seal. If silicone is added to the outside of the enclosure it should be black and must not impeded the sensor making contact.

17. Cut, strip & tin cable assembly wires to approximately the correct lengths . These can be shortened as production technique improves.

18. Solder R[image: image6.jpg]


ed & Blue wires to battery terminals. This is best done  first

1. Red to Right looking into enclosure right way up[image: image7.jpg]Long cntr

YELLOW

Left

Violet




2. Blue to left

3. Batteries  only work one way.

19. Solder Green to centre of Jack and Brown to outside.

20. Solder wires to switch (Black & White)  

21. Solder wires to the pot see Diagram 

1. Orange Right

2. Yellow Centre

3. Violet Left 


22. Add Battery labels 

1. +ve Inward  Right  -ve on the visible edge

2. -ve  Inward left. +ve on the visible edge. 

23. Using test battery ( Two batteries shorted at one end).

24. Check the unit functions using a test LCD/PCB 

25. Installation of PCB & LCD

26. The excess pips on the side of the PCB may need removing

27. NB  If the PCB has a serial number. This must be recorded for traceability. If not record the batch number.

28. Plug the LCD lead into the PCB socket

29. Connect the LCD to the PCB with a small amount of double-sided tape e.g 2.0 cm x 1.0cm (this allows removal without damaging the LCD)

30. Take care the LCD input cable is very fragile. It breaks just after the connector

31. Take care not to bend the ribbon cable too sharply

32. Add two small amounts of Blu-Tack (or equivalent )one on either corner of the LCD mounting. This is to help alignment of the PCB in the front panel.

33. Introduce the PCB into the enclosure. 

34. The connectors should be at the top of the instrument..

35. Ensure alignment is good and the LCD lies level. The Blu-Tack allows the LCD to be aligned & levelled correctly

36. Remove the LCD protective cover

37. Add an extra Polycarbonate lens

38. Locate with Blu-Tac or equivalent.

39. Ensure the faces of the LCD & both faces of the extra lens are dust free etc.

40. Locate the seal on the back of the front panel ensure it fits under the locking lugs

41. Offer the front panel up to the enclosure 

42. It locates with small fragile lugs on each side and a lock on the top.

43. Ensure LCD alignment

44. The front panel can be removed without damage if very great care is taken. The top should be released first. Then the bottoms gently prised over the plastic hooks.

45. Screw the front panel on with 2 size M2x10 screws from the rear. 

46. [image: image8.jpg]9530003

9530012
9530013

9530006
9530004 ; ' s
u f i | 9530002

9530006

- 9730061



Add the two outer side covers and fix with M2 x 5 screws. Take care not to over tighten 

47. Add serial number label  White to inside 

48. Add same serial number  Black Label to rear of Instrument

49. Using Test Battery and 

50. Using Test box  or mV source 100% = 50mV Air =21%  10.5 mV approx. +/- 1%

51. Add Sensor ; note serial number[image: image9.jpg]Green 90mm  Jack centre '\, _
Brown 90mm Jackgnd Black 80mm Switch
White 80 mm Switch

Orange 80mm Potright
Yellow 80mm Pot cntr

Violet 80mm Pot Left

Red 90mm Batt +ve
Blue 90mmBatt -ve

Cable sub assembly PA 29/11/2007




52. Rotate to ensure no contamination on connectors

53. Add batteries and fix cover with three M2 x 10 screws

54. Check the unit functions

55. Add the front labels 

56. Switch On

57. Check correct software version

58. Check correct title[image: image10.jpg]



59. Check the correct colour switch 

60. Yellow auto switch off .Green no auto switch off

61. Set Calibration knob to read  20.9%

62. Add labels to box

63. Add Quick-OX ;

64. Leave the Flow divertor attached

65. NB DINKIT needs to be supplied separately

66. Add user guides[image: image11.jpg]




Appendix A

Software:

The current main software is :-

Version 4.0

Legend;  Vandagraph TEK-OX  alternating

Auto switch-off

Auto-switch off can be disabled by breaking the shorting link on the PCB

The software can be changed to Derived Nitrogen by breaking a link on the PCB

By using the Logo maker programme the Legends can be changed to two 10 letter words. More letters will reduce the size and the quality of the print  

Carrying case

An alternative case is the Otter Extreme 9000.

The Quick-ox fits in the lid and the analyser in th main part.

Labels should be fitted 

A User sheet should be placed in the box


ERRORS:

E1: LCD socket not 100% aligned or faulty LCD

No Reading: Remove battery to reset software ( mainly when using test unit)



Check batteries are in correct. Faulty LCD, bad battery contact

OO  With sensor: Rotate sensor to ensure Jack plug is clean

Equipment Required:


Soldering Iron: side cutters: Spanner;Long nose pliers. Tweezers, solder tools, carver set











VM3COP42.01 TEK-OX Assembly 080811.doc         04/13/08                                      Page 5 of 5

