Frequently Asked Questions   Helium Analysers


Questions
Answers

How long to warm up   
seconds

Do I need a Flow meter
NO 

1. A fixed 2 litres /min flow restrictor DIN fitting  is supplied. Open until a gentle hiss is heard add the tubing and wait until the reading stops climbing. Take reading Switch off the cylinder the both O2 & He readings may quickly fall to back pressure. This is the reading you require

How do I calibrate 
1. Oxygen  20.9% in air

2. Oxygen 100& in O2

3. Helium set zero in air

4. Use 100% Helium to set the span

How do I check accuracy
1. Use 100% Oxygen and 100% Helium and mix through a Tee piece. 

2. Adding the O2 reading to the He reading should be about 100% ( usually 101%



Why? When I check 100% Oxygen the He reads Approx 1.7%
Oxygen Has a higher Thermal conductivity than air so reads higher.

I cannot set the O2 sensor down to 20.9%
Oxygen is difficult to remove from the sensor and can take several minutes,

Flush out the system with air using the restrictor 

I get a different reading in pen air to air from a cylinder
Air from the cylinder is dry. Fresh air contains moisture and therefore reads a lower level of oxygen

How do I know the instrument is correct
Feed 100% Oxygen and 100% Helium into the instrument. When added together the reading should be around 100% +/- 2% ( errors due to flow)

What is life of Oxygen sensor
Estimated 36 months in Air.

This is based on a 70% anode consumption as 100% is impossible.



What is the life of the Helium sensor
7 to 20 years if looked after

H:\Archive Vandagraph 080324\Heliun Analysers\FAQ Heilium Mixcheq.doc
7/17/2008

