Procedures for the Construction of


VN202 Oxygen Analyser





Ref�
Description of action�
Time�
�
�
Preparation of the Case�
�
A.1.0�
Unwrap the new cases and separate them into their component parts, be careful not to lose the case screw which is inside the battery compartment, and the window which is inside the case.�
�
�
�
A.2.0�
With a strong, sharp blade cut the two mounting pegs nearest the top on the back case down by approximately 1-2mm.�
�
�
�
Screening the cases�
�
�
�
B.1.0�
Wrap masking tape around the sides of the back case, so that about 2mm of tape can be folded into the inside edge of the groove that runs around the case.�
�
�
�
B.2.0�
Cut away the piece of tape covering the bottom of the case which forms part of the battery compartment.�
�
�
�
B.3.0�
Wrap tape around the front case, covering the outside but not the inside of the ridge which runs around it.�
�
�
�
B.4.0�
Wrap tape around the edge of the battery door.�
�
�
�
B.5.0�
Place all the masked parts down on a suitable surface with their insides upwards for spraying.�
�
�
�
B.6.0�
Place the top panels down with their largest face upwards as this will become the inside surface.�
�
�
�
B.7.0�
With the screening spray, cover the exposed surfaces from every angle. Allow approximately 15 minutes to dry then repeat.�
�
�
�
B.8.0�
Remove tape.�
�
�
�
Drilling the cases�
�
�
�
C.1.0�
Stick the top labels to the unsprayed surface of the top panel.�
�
�
�
C.2.0�
With a 6mm drill bit drill a hole between the Vandagraph logo and the address, below the centre of the panel.�
�
�
�
C.3.0�
With a lOmm bore, drill a hole in the front case exactly in the middle from left to right and l4mm from the top of the label area.�
�
�
�
C.4.0�
With a l2mm bore drill a hole in the front case below the first one, exactly in the middle from left to right and l4mm from the bottom of the label area�
�
�
�
Final preparation of case�
�
�
�
D.1.0�
remove any dust from the front case and carefully stick on a front label, ensuring there are no bubbles or wrinkles.�
�
�
�
D.2.0�
From the inside of the case, cut around the holes with a scalpel blade.�
�
�
�
D.3.0�
On the larger hole, cut a triangular notch about 2mm deep.�
�
�
�
D.4.0�
Put a switch seal onto the switch and put the switch into the front case aligning the tag with the notch.�
�
�
�
D.50�
Place the small bar onto the back of the switch and screw it in.�
�
�
�
D.6.0�
Put the rubber washer onto the potentiometer and put into the front case.�
�
�
�
D.7.0�
Put the crimped washer and the nut onto the potentiometer and tighten, ensuring that when viewed from the front the white dot on the knob moves equal distances clockwise and anti-clockwise from the middle position.�
�
�
�
D.8.0�
Squirt a small amount of silicone rubber into the channel moulded into the front case and smooth it flat.�
�
�
�
D.9.0�
Put a battery door into a case back, then stick masking tape over the area where the battery door catch can be seen.�
�
�
�
D.10.0�
Cover the tape with silicone rubber and fill along the bottom edge of the back case to ensure a watertight seal.�
�
�
�
D.11.�
Push a grommet through the top panel from the front making a neat fit if necessary by widening the hole slightly with a reaming tool.�
�
�
�
Printed Circuit Board�
�
�
�
Assembly�
The PCB requires assembling using the principle that components of the lowest height are placed in first. Polarity of IC's, the diode Dl and the capacitor C2 must be observed�
10�
�
�
E.1.0�
Insert link wire L, ensuring it is flat to the PCB.�
�
�
�
E.2.0�
Insert all resistors.�
�
�
�
E.3.0�
Insert diode observing correct polarity.�
�
�
�
E.4.0�
Insert IC 1 and IC 2 ensuring they are correctly aligned by observing the markings on the PCB.�
�
�
�
E.5.0�
Insert the 28-way socket.�
�
�
�
E.6.0�
Insert Cl and C3, polarity is unimportant.�
�
�
�
E.7.0�
Insert C2 noting that it must be inserted with the negatively marked leg to the outside of the PCB so as to connect to the negative track on the other side of the board.�
�
�
�
E.8.0�
Insert FET 1 so that the flat side is to the outside of the PCB.�
�
�
�
E.9.0�
Insert P1 with the adjustment screw to the right hand side.�
�
�
�
E.10.0�
Insert the battery connector ensuring red is to + and black to -.�
�
�
�
E.11.0�
Take 5 pieces of wire approximately 7cm long, strip and tin both ends.�
3�
�
�
E.12.0�
Solder two pieces into the PCB at the position marked SW.�
�
�
�
E.13.0�
Solder the other three pieces into the PCB at the position marked VR1.�
�
�
�
Spraying the PCB's.�
�
�
�
F.1.0�
The PCBs require spraying with a conformal coating.


Always use in a well ventilated area.


Always follow the instructions on the can.�
�
�
�
F.2.0�
Stick a piece of masking tape over the 28-way socket.�
�
�
�
F.3.0�
Take a piece of copper tape about 6cm long and peel back 1cm of the backing, stick this to the back of the PCB on the exposed part at the bottom at right angles to the bottom edge of the PCB.�
�
�
�
F.4.0�
Spray the back of the beard with conformal coating. Allow to dry. Turn the PCB over and spray the front.�
�
�
�
Attaching the Coiled Cable�
�
�
�
G.1.0�
Fold dual plugged cable in half and cut into two equal lengths.�
6.5�
�
�
G.2.0�
Strip about 4cm of outer sleeving being careful not to damage the shield or the inner conductors.�
0�
�
�
G.3.0�
Separate the shield from the wires, twist It and shrink about ~m of narrow heat shrink tubing over the shield.�
0�
�
�
G.4.0�
Strip about 3-4mm of insulation from the inner conductors.�
0�
�
�
G.5.0�
Carefully tin the Inner conductors.�
0�
�
�
G.6.0�
Push the cable through the grommet of a pre-prepared end panel.�
0�
�
�
G.7.0�
Remove the masking tape from the 28-way connector on the PCB.�
0�
�
�
G.8.0�
Insert a tie-wrap down through the bottom of the two holes under the 28-way connector, loop it back through the top hole and through the locking part of the tie-wrap.�
0�
�
�
G.9.0�
Push the coiled cable through the tie-wrap loop on the underside of the board so that the inner conductors are nearest the solder pads where they are to be connected, allow the tie-wrap to grip the cable about 5mm from the end of the sleeving then with a pair of pliers pull the tie-wrap tight. Snip the excess from the tie-wrap.�
0�
�
�
G.10.0�
Solder the yellow wire into the top solder pad and the grey wire into the bottom solder pad ensuring that there isn't any strain on the wires.�
0�
�
�
G.11.0�
With a soldering iron, on the back of a panel meter join the two halves of the solder pad marked "2".�
�
�
�
G.12.0�
Insert into 28-way connector ensuring that the side of the DPM with only 10 pins is nearest the top of the board.�
�
�
�
Testing the PCB�
�
�
�
H.1.0�
Solder the 3 wires from the beard at the point marked "VRl" to the potentiometer on the front case. They must be connected so that the wire nearest the top of the beard goes to the far tag on the pot, the middle wire to the middle tag, and the bottom wire to the nearest tag.�
�
�
�
H.2.0�
Solder the 2 wires from the board at the point marked "SW" to the switch on the front case.�
�
�
�
H.3.0�
Connect the battery connector to a variable power supply set at  9V.�
�
�
�
H.4.0�
Press the switch on the front case to turn the unit on.�
�
�
�
H.5.0�
Connect the jack plug to a millivolt source, switched off so as to give zero output.�
�
�
�
H.6.0�
Set the pot on the front case to fully clockwise.�
�
�
�
H.7.0�
Adjust the trimpot P1 until a reading of 00.0 is observed. Set the millivolt source to 5OmV.�
�
�
�
H.8.0�
Adjust the pot on the front case to calibrate the monitor to 100.0, the reading should not drift.�
�
�
�
H.9.0�
Adjust the millivolt source until the LC D reads 88.8 to check the integrity of the LCD segments.�
�
�
�
H.10.0�
Adjust the millivolt source down to l0.5mV, the LCD should read 21.0 + 0.2�
�
�
�
H.11.0�
Switch the millivolt source off checking that the display reads 00.0 adjusting the trimpot P1 if necessary.�
�
�
�
H.12.0�
Reduce the voltage from the power supply until the low battery indicator can just be seen to appear, the display should not drift by more than 1.0%.�
�
�
�
H.13.0�
Turn the unit of and back on again. The time elapsed before the auto shut off is activated should be approximately 3 minutes.�
�
�
�
Encapsulating the P.C.B.�
�
�
�
I.1.0�
Cut a strip of conductive foil tape approximately 6cm long�
�
�
�
I.2.0�
Stick one end to the back case on the bottom left hand side about 1cm up from the mounting peg, ensure that it is firmly stuck down the inside of the casing and at least 1cm across the back.�
�
�
�
I.3.0�
Stick the other end to the front case in the same way.�
�
�
�
I.4.0�
Peel the backing off the tape on the PCB and stick this to the back case above the silicone battery door seal.�
�
�
�
I.5.0�
Screw the PCB to the back, leaving the screw nearest the battery connector until last.�
�
�
�
I.6.0�
Put the battery door on the back of the case and connect a battery. Place the battery in the battery door so that the connector is on the right hand side of the case.�
�
�
�
I.7.0�
Pull the battery wires tight and tuck them into the channel moulded into the case.�
�
�
�
I.8.0�
Put a washer onto the last screw, separate the red and black wires so that they are to either side of the screw hole and insert the last screw, be careful not to over tighten�
�
�
�
I.9.0�
Pull the down the coiled cable and align it so that it is ready to go into the back case i.e. so that the grommet is to the front.�
�
�
�
1.10.0�
Pull the shield wire from behind the board so that it will be accessible from the front when the top panel is inserted.�
�
�
�
I.11.0�
Take a syringe filled with black silicone sealant and fill the gap between the PCB and the back case, including the groove in the back case that the top panel fits into.�
�
�
�
I.12.0�
Fill between the DPM and the beard.�
�
�
�
I.13.0�
Insert the top panel into the back case, ensuring it goes right into the groove in the back case, then pull the cable through the grommet as far as possible�
�
�
�
I.14.0�
Take about 3cm of copper tape and stick the splayed shield to the top panel near the left hand case tag, stick the other end of the tape to the tag.�
�
�
�
I.15.0�
Fill the gap between the DPM and the top panel, being thorough around the grommet area.�
�
�
�
I.16.0�
Fill around the DPM and put a bezel over it. It may be necessary to pull on the coiled cable to allow the bezel space to fit.�
�
�
�
I.17.0�
Fill the gap around every side of the PCB,�
�
�
�
I.18.0�
Allow the sealant to cure this is best done overnight.�
�
�
�
Sealing the instrument�
�
�
�
J.1.0�
Retest the instrument as before, adjusting the trimpot back to 00.0 if necessary.�
�
�
�
J.2.0�
With a multimeter test the continuity between the back case, front case and top panel. if the meter can be made to ring through then the continuity is good.�
�
�
�
J.3.0�
Take a DPM window and peel off any stickers on the front but not the protective film.�
�
�
�
J.4.0�
Peel off the card on the back of the window to expose the adhesive, peel off the protective film that is on the DPM itself and stick the window to the DPM being careful not to trap any dust between the two.�
�
�
�
J.5.0�
Partly close the case to see if the window looks level, this is best done with the instrument on so as to see if the display looks correct. Adjust the window and retry if necessary.�
�
�
�
J.6.0�
Take a case window and insert it firmly into the front case, this is best done by using the plastic bag that the window came in to push the window into place to avoid creating fingerprints�
�
�
�
J.7.0�
Gently run a soldering iron around the edges of the window to fix it into place.�
�
�
�
J.8.0�
Take the front case and hold it over the top of the back then run a line of silicone sealant down the groove on the left hand side of the back case.�
�
�
�
J.,9.0�
Carefully lift the front case back over to the left hand side trying not to touch the line of sealant just laid.�
�
�
�
J.10.0�
Lay sealant along the edge of the top panel, the bottom and right hand edges of the back case.�
�
�
�
J.11.0�
Peel the protective label off the DPM and quickly close the case ensuring not to trap any wires particularly at the bottom edge. if wires do get trapped they can be pushed in with a thin screwdriver.�
�
�
�
J.12.0�
Insert a case screw and tighten. Do not over tighten.�
�
�
�
J.13.0�
Place the instrument into a plastic hobby vice which has its jaws covered with foam rubber. It should grip the instrument about 1cm from the top and not be too tight.�
�
�
�
J.14.0�
Make one final check that the edges of the case are snapped closed as far as they can go before leaving the sealant to cure, again leave overnight if possible.�
�
�
�
Finishing off the analyser�
�
�
�
K.1.0�
Retest analyser as mentioned in Testing the PCB.�
�
�
�
K.2.0�
Remove all the excess silicone rubber from the case, including inside the battery compartment.�
�
�
�
K.3.0�
Any rubber in the side grooves of the casing can be removed with a scalpel.�
�
�
�
K.4.0�
Stick a back label onto the case.�
�
�
�
K.5.0�
Stick a label on the front at the bottom.�
�
�
�
K.6.0�
These two labels are the only two labels that will be different if the monitors have been made with customer specific labels.�
�
�
�
K.7.0�
Stick a serial number label inside the battery compartment.�
�
�
�
K.8.0�
Place battery in the compartment and close the battery door.�
�
�
�
K.9.0�
Ensure the cable is clean and free from rubber.�
�
�
�
K.10.0�
Clean the entire instrument with silicone polish using a soft cloth or kitchen roll.�
�
�
�
K.11.0�
Place the analyser into a bag making sure the coiled cable is down the side of the instrument and not across the labels as this can leave a lasting mark on them.�
�
�
�
K.12.0�
Place the instrument into a plastic hobby vice which has its jaws covered with foam rubber. It should grip the instrument about 1cm from the top and not be too tight.�
�
�
�
K.13.0�
Make one final check that the edges of the case are snapped closed as far as they can go before leaving the sealant to cure, again leave overnight if possible.�
�
�
�
K..14.0�
Retest analyser as mentioned in Testing the PCB.�
�
�
�
K.15.0�
Remove all the excess silicone rubber from the case, including inside the battery compartment.�
�
�
�
K.16.0�
Any rubber in the side grooves of the casing can be removed with a scalpel.�
�
�
�
K.17.0�
Stick a back label onto the case.�
�
�
�
K.18.0�
Stick a label on the front at the bottom.�
�
�
�
K.19.0�
These two labels are the only two labels that will be different if the monitors have been made with customer specific labels.�
�
�
�
K.20.0�
Stick a serial number label inside the battery compartment.�
�
�
�
K.21.0�
Place battery in the compartment and close the battery door.�
�
�
�
K.22.0�
Ensure the cable is clean and free from rubber.�
�
�
�
K.23.0�
Clean the entire instrument with silicone polish using a soft cloth or kitchen roll.�
�
�
�
K.24.0�
Place the analyser into a bag making sure the coiled cable is down the side of the instrument and not across the labels as this can leave a lasting mark on them.�
�
�
�



z:\main\cemark\vn202\vnconstr	Page�page  \* MERGEFORMAT�1�    		5 October 1998











