Blu-Ox  Temperature tests


The Blu-Ox specification is 0oC to 30oC

This text is to take the monitor below 0oC   I.e -3oC   for 7 days

The Blu-Ox was placed into the freezer compartment of a fridge

The temperature was measured  using a Fluke 867B multimeter ( S/N 6911002 Calibrated)

and an 80TK thermocouple module 

A DC voltage was inserted into thre Blu-Ox with a Microcal set to 20.0mv.

The Cal control was set to read 49.9% and fixed. At 19 oC

26/09/08  slow switch on       49.9%

29/09/08    “        “                 49.9%   ( 3 measurements throughout the day)

30/09/08    “        “                 49.9%

03/10/08    “        “                 49.8%      

06/10/08    “       “                   49.7% 

J.S.Lamb

06/10/08

10/06/08   Temperature Tests


