There has been a report that the Microstim can give erroneous signals.





This has been confirmed if the switches are switched from one mode to another mode very quickly.





This is due to the mechanical inertia in the switch.





The switches have been used since 1986 without any other reported problems.


There is a possibility that no new switch could be substituted to solve the effect.





The instruction manual clearly states that the switch should be allowed to return to rest ( I.e off) before another action is carried out. This statement will be re-inforced in future manuals.





There are no clinical procedures that require the fast movement from one state to another unless the clinician has chosen the wrong mode.





The only effect on the patient could be one of discomfort.








J.S.Lamb








10/10/2008








 


Risk Report Microstim switches





Risk Report Microstim switches                     10/10/2008











