[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid






[image: image2.png]


 Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire, BD20 7DT, UK.

Website: www.viamed.co.uk.  Email: info@viamed.co.uk.

Tel: +44 (0)1535 634542.  Fax: +44 (0)1535 635582.

.

[image: image1.jpg] Tom Thumb Resuscitation Unit Servicing Manual (TT490).

Version 1.1.  16/07/04



Part no: 0390023.


TOM THUMB CALIBRATION / SERVICE SHEET.

Description.
Tom Thumb.
Serial No.


Model.
TT 480.
Time & Date of Service / Calibration.


Record the manometer reading in millibars below, ensuring the TT480 meets the limits specified.

Test Equipment
Test
Specification
Reading
P/F




Calibration








Snoop liquid.
Check all ports / connections for leaks.
No bubbling.
√
P







Manometer
Adjustable Valve:

@ 10 Lpm
Minimum:

 8 cmH2O.

p







Manometer
Adjustable Valve:

@ 10 Lpm.
Tolerance:

 42 &  47 cmH2O.

p







Manometer
Precision Valve:

@ 10 Lpm.
Maximum: 50 cmH20

Minimum 2.0 cmH2O Higher than the adjustable valve setting.

p







Manometer
Pressure Gauge Test:

@ 30 cmH2O.  
Tolerance:                                 29 &  31 cmHO

P







Visual check.
Gauge cannot be removed without tools.

√
P







Visual check.


Gauge appears straight.

√
P







Visual check.
All adjustable settings are set to a minimum.

√
P







Visual check.
Labels are attached.
CE Label.
√
P



Serial No. Label.
√




Tom Thumb Label
√


Service.
Parts Replaced.

Changed “O” Rings

Diamond Copper Washer


Re-greased Adjustable Valve Internal Parts

Adjustable Valve


Re-calibrated Tom Thumb

Adjustable Valve Internal Parts




Safety Valve




Safety Valve Internal Parts




Pressure gauge




Pressure Gauge Face




“O” Rings




Outlet




Inlet 




Blanking Bolts




Rail Clamp




Main Body Block


Calibration: …E. Sendrowski ………………………………
Signed:  …………………………………           

[image: image2.png][image: image3.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




