Rationale for Vandagraph EOII sensors


There are two Analox sensors for the EOII analyser. Unfortunately both have the same part number although a Reference Number may be different.

Analox part number is 9100-9220-9

[image: image1.jpg]



One uses a small JST socket and has a high output 25mV+/- 2mV.  In Air

Analox offer a kit SA2EII9B

The Vandagraph equivalent is the R-33DE.

This is constructed using VM3COP40.13 R-33DE sensor assembly

[image: image2.jpg]


The other 4.5 mV+/- 1mV  7mV- 11mV some very  high output sensors may not calibrate down to 21% due to analyser tolerances so far unproven

The Vandagraph equivalent is the R-22DEM

This is constructed using VM3COP40.14 R-22DEM sensor assembly

The outputs required to drive the analysers has been checked with a Microcal

Four analysers have been tested. The first was tested with the Microcal open circuit.

And was tested over a range of readings. Serial number unknown,

A Further three were tested one serial number unknown.

In order to be able to match a Teledyne sensor one analyser was tested with a 4k7 and then a 1k resistor across the Microcal output. The difference was negligible so a 4K7 (or 5K6) will be used in any further tests. 4K7 was chosen as the output of an R-22 can be halved using 2 off 4K7 resistors giving the R-22 the ideal input impedance of 10K.

The analyser tested without an output resistor shows a slightly higher curve but still within the tolerance of the sensor output ranges.

[image: image3.jpg]70

60

50

40

30

20

10

60

50

40

30

20

10

Title

Low set
fi—
High Set

——

4.5
4.8
5.1
5.4
5.7
6.6
6.9
7.2
7.5
7.8
8.1
9.3
9.6
9.9
10.2
10.5
10.8

Title

Ell Lo IK olp
—{E—

Ell Lo 4.7K olp
—_——

Ell lo S/N 529
——

EllHi 1K olp
ey

Ell Hi 4.7K o/p
——

EIl Hi S/N 529
g

1 2 3 4 5 6 7
Oxc voltage

0




EOII Vandagraph Rationale        16/10/2007         Page 1 of 1

