Vandagraph

Replacement of the Oxyspy Oxygen sensor

Using 

Vandagraph  Teledyne R-22D

The following notes are a guide

[image: image1.jpg]



1. Refer to manufacturers instructions. 

2. Pull the yellow ring off the rear of the case

3. Loosen the 7 screws located around the case.

4. Carefully separate the two parts.

5. Disconnect the connector carefully from the back of the sensor

6. [image: image2.jpg]


Remove the sensor from the unit.

7. Remove the new sensor from its pack and check it for leaks

8. Remove the new sensor 0-ring .

9. Attach the yellow mounting ring from the old sensor onto the new sensor by gently twisting until it fits snugly

10. [image: image3.jpg]


Ensure that the prongs on the sensor  and sockets on the connectors line up and push in firmly.

11. Fit the sensor with the yellow mounting ring back into the housing so that the ring fits into the notch.

12. Ensure the wires are not trapped under the sensor 

13. Replace the cover carefully taking care that the sensor locates properly and that you do not trap any wires

14. [image: image4.jpg]


Do not over tighten the screws.

15. Refit the rear yellow ring ensuring that gas can get to the sensor

16. Leave exposed to dry outdoor air for at least two minutes NB (All galvanic sensors can take an hour to fully stabilize after they have first been exposed to air to achieve maximum accuracy. This may be visible as a rising output)

17. Calibrate to 20.9%

18. [image: image5.jpg]


Dispose of the old sensor according to local regulations for Lead and Potassium Hydroxide solution

[image: image6.jpg]



Vandagraph Ltd.  15,Station Road, Crosshills , Keighley, West Yorkshire. BD20 7DT

+44 (0) 1535 634900   Fax: +44 (0) 1535 635582

www.vandagraph.co.uk      email  sales@vandagraph.co.uk
Oxyspy  sensor replacement 070307.doc
Page 1
07/03/2007

