OPERATING PROCEDURE

COP40.09

Assembly kit comprises

[image: image10.jpg]



1. TUBING VP12   
Stock no
9710040

2. DIN22F
Stock no 
9711006

3. Serial number label
Stock no

4. Cannister
Stock no


5. A268


Stock no
9711001 

6. D22M
 
Stock no
9711004

7. O RING   BS111
Stock no
9711003

8. Canister Label
Stock no

[image: image2.wmf]
[image: image3.wmf]
Insert the O ring into the groove

NB do not use any grease as the product must be Oxygen clean

[image: image4.wmf]
[image: image5.wmf]
[image: image1.wmf]
Remove the R-33D from the gas barrier bag and carefully insert it into the holder pt no,  so that the wires can be bent to meet one of the two grooves    Ensure the sensor is pushed firmly home Do not damage the PCB components or wires

[image: image6.jpg]o





[image: image7.jpg]


[image: image8.jpg]



Write the serial number of the R-33D on twp labels 

Small legible writing is required

[image: image9.jpg]


Carefully remove paper backing and stick transparent cover down.

Carefully mount one label on the sensor & 

one label on the white container so that the edges

of the labels are parallel to the edges of the container and sensor .

Test the sensor    Output should be   23 mV   to  27  mV

Place packing in bottom and

top of the white container and add cap.

Ensure it is sealed.

Add Vandagraph round label to top

Add Instruction leaflet 

Place in a polythene bag and seal







ISSUE: 1
REVISION: 1
10 June 2003

W:\COMPANY\Iso9000\ISO VM3COP v2002\VM3COP40 Vandagraph\VM3COP40.09    VN33Van assembly2.doc

