	
	
	
	

	
	
	
	
	
	Pre-use Checks.

· Uncap the T piece port (.
· Adjust the flowmeter ( to minimum (fully clockwise) and the adjustable valve control ( to minimum (fully counter clockwise).
· Connect the flowmeter O2 hose ( to the external O2 supply.
· Check that the pressure gauge ( reads zero. 

If not, the Tom Thumb requires servicing.

· Connect the patient tubing ( to the Tom Thumb outlet but do not apply the mask ( to the patient.

· Set the flowmeter to the required flow rate.

· Occlude the mask and the T piece port. 

Gradually turn the adjustable valve control clockwise until the required outlet pressure is shown on the pressure gauge.
· The Tom Thumb is now ready for use.

[image: image1.png]



Important.

For use by qualified and trained personnel only.

Use flow rates within the range of the flowmeter.

Adjust outlet pressure after altering the flow rate.

Do not attempt to adjust the safety valve (.

Recommended O2 inlet pressure of 4 bar.


	
	

	
	
	[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





	[image: image3.png]



	PN 0390007.

V2.1.
	
	
	
	
	
	

	
	
	25/10/04

Tom Thumb, TT490-15

Instructions for Use

[image: image4.jpg]



	
	
	
	
	
	

	
	
	
	
	
	

	
	
	
	

	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	Guideline for Use during Resuscitation.

1. Follow the pre-use checks and set the required flow rate and outlet pressure.

2. Apply the mask to the patient and cover the T piece port to inflate the patients’ lungs at the set flow rate and pressure.

3. Uncover the T piece port and allow the patients lungs to deflate.

4. Repeat steps 2 & 3 as necessary during the resuscitation of the patient (follow the hospital protocol for resuscitation).

Care, Cleaning, Location and Sterilisation.  

Clean using a damp cloth. The Tom Thumb is not intended to be sterilised. Do not autoclave. Do not allow moisture or foreign matter to enter the safety valve or adjustable valve.  Damage will occur if the Tom Thumb is subjected to severe mechanical shock or dropped.  

The Tom Thumb should be serviced every 12 months, if the pressure gauge does not read zero (outside of the black band) with no flow or if the units’ accuracy is doubted.

The Oxygen Hoses should be checked every 3 months, and replaced ever 4 years as a minimum.

The rail bracket ( is designed to fit most medical rails.  It is advised that the Tom Thumb is not mounted close to a wall or to the side of an incubator particularly if the gauge is fitted to the end of the body (specials only).  
	
	
	
	
	Warranty.

Viamed warranty ensures that goods are free from defects of manufacture for a period of one year from the date of shipment from Viamed. 

Liability shall be limited solely to the replacement and repair of the goods and shall not include shipping costs or other incidental damages.

This warranty is null and void if any items are subjected to misuse, negligence, accident, or repairs other than those performed by Viamed or an authorised service centre.

Viamed Ltd.

15 Station Road,

Cross Hills,

Keighley,

West Yorkshire,

BD2O 7DT,

United Kingdom.

Tel: +44 (0)1535 634542.


Fax: +44 (0)1535 635582.

                   Email: info@viamed.co.uk


                   Website: www.viamed.co.uk
© Copyright 2003. Viamed Ltd. All Rights Reserved
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	


























9



























































8





4





6





3





7





2





5





1









