[image: image2.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR



                                                                                                                 

Vandagraph Limited  - 15 Station Road  -  Cross Hills

           

Keighley  -  West Yorkshire BD20 7DT  -  United Kingdom

           

Tel +44 1535 634900                       Fax +44 1535 635582   

            
Email: info@vandagraph.co.uk     www.vandagraph.co.uk


Specification Datasheet - 8010060

WITT-Gasetechnik GmbH 

OSGA-20 Oxygen Sensor
	Specifications


	Vandagraph part number
	8010060

	OEM model number
	OSGA-20 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1- hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	> 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	> 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



v1.0

21st January 2021
Page 1 of 3
	Drawings/images


[image: image3.png]WITT-Gasetechnik GmbH & Co KG Tel +49 (0)2302 8901-0
Salinger Feld 4-8 Fax +49 (0)2302 8901-3
58454 Witten witt@wittgas.com
Deutschland/Germany www.wittgas.com

OSGA-20 Oxygen Sensor

M) 202011 (1 (€
SIN W123456 %

(LR 2

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes and
skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from this bag
until ready to use. Depending upon the ambient conditions the oXygen Sensor may require
30 minutes to reach signal stability after installation. )





All dimensions in mm

Page 2 of 3

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image1.png]0OSGA-20 Oxygen Sensor

SIN  W123456

Tomme /N 98 C€ 202011 T





Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image4.jpg]({) VANDAGRAPH LTD





Page 3 of 3


