[image: image1.jpg]


 Tom Thumb Resuscitation Unit Servicing Manual (TT490-15).
Version 1.7. 07/04/06

Part no: 0390023. 

[image: image30.wmf]Tom Thumb Resuscitation Unit Servicing Manual (TT490-15).

Version 1.1. 16/07/04.

Part no.: 0390023. 


	TOM THUMB RESUSCITATION UNIT 

TT 490-15

SERVICING MANUAL




[image: image31.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid







INDEX
Section
 






          
                                    Page













1.  Introduction










  2

2.  Servicing: TT490-15








  
  3

3.  Calibration Sheet









  8

 ----------------------------------------------------------------------------------------------------------------------------------

This manual is intended to provide information to help qualified maintenance personnel service and repair the Tom Thumb Infant Resuscitation Unit; TT490-15. General engineering knowledge and the ability to follow technical instructions are assumed, as are knowledge of oxygen flow rates and the characteristics of operational pressures.

The equipment needed to service the Tom Thumb is laid down in this Service manual. In addition, a calibrated manometer will be required for test and calibration of the unit.

Diagrammatic representations of disassembly and re-assembly are shown in this Service manual.

Servicing personnel must be aware of the potential clinical implications of incorrectly serviced equipment.

Service and calibration of units, to a flow of 15 Lpm, does not imply operational usage at this rate. Operational flow rate should be determined, at the time of usage, within the scope of individual organisations’ resuscitation protocols, up to the maximum of 15 Lpm.

-----------------------------------------------------------------------------------------------------------------------------------

	1.  Introduction.


Service of the Tom Thumb Resuscitation Unit.

The Tom Thumb has been designed to require minimal service with very few replaceable items. The accuracy of the pressure gauge, adjustable valve and precision valve should be checked at least every 12 months or when the gauge at zero pressure reads outside the black band.

It is recommended that all hoses be checked every 3 months, for connections, damage and possible degrading and / or rupture of hose and coatings. The hoses should be replaced, regardless of condition, every 4 years (maximum).

It is recommended that all O-rings should be replaced every year.

Setting of the adjustable valve pressure when carried out in accordance with the relevant procedure will ensure accuracy to ± 2 cmH20.

The precision valve pressure may have been set at manufacture to 20, 30, 40 & 50 cmH20 to a tolerance of ± 2 cmH20, dependent on customer requirements.  Setting of the precision valve when carried out in accordance with the relevant procedure will ensure accuracy to ± 2 cmH20.  

If adjustment/set up of adjustable or precision valves is required, valves must be removed and serviced according to Service Manual for adjustable or precision valves. Alternatively, instruments can be returned to Viamed for servicing as specialist tooling is required to dismantle and reset.

The Tom Thumb pressure gauge is accurate to ± 6 cmH20 subject to a reading within the black band at zero pressure.  If the gauge is removed, a single-use diamond copper washer, part no. 0330212, must be replaced.  

[image: image2.png]



	2. 0310032 TT490-15 / 0310034 TT490-15 (1m & 3m hoses) Servicing.


Important: Use only oxygen-compatible grease during assembly of Tom Thumbs.  Do not use or allow organic greases to enter the Tom Thumb.  Ensure all parts are clean before assembly.

	Equipment required.
	12mm open-ended spanner, pick, isopropyl alcohol, kitchen tissue, adjustable spanner, PTFE tape, 3mm Allen key, 7/8 A/F spanner, 5/8 – 9/16 A/F spanner, O-Rings, diamond copper washer.


	Parts list.

	Qty.
	Description.
	Part No.

	If req’d
	Body block
	0330201

	1
	Diamond copper washer
	0330212

	If req’d
	Pressure gauge
	0330203

	1
	Thick O-ring – 3/32 section
	0330213

	6
	Thin O-ring – 1/16 section
	0330214

	If req’d
	Flowmeter bolt
	0330205

	If req’d
	Flowmeter
	0320060

	If req’d
	Precision valve
	0330210

	If req’d
	Blanking bolt
	0330207

	If req’d
	Adjustable valve
	0330211

	If req’d
	15mm outlet
	0330209  

	If req’d
	Serial number label
	0390032 or 0390034

	If req’d
	Inlet hose right angle adapter
	0330219

	If req’d
	Inlet hose (1m)

Inlet hose (3m)
	0330217 or

0330218

	If req’d
	Spacer block
	0330215

	If req’d
	Rail Clamp
	0330055

	If req’d
	M4x10mm hex drive bolts
	0330216

	If req’d
	“Tom Thumb” label
	0390015

	If req’d
	Flowmeter label
	0390012

	If req’d
	Adjustable valve label
	0390010

	as req’d
	Oxygen-compatible grease
	0330220

	as req’d
	Oxygen-compatible adhesive
	0330237


	Servicing.


1. Detach rail clamp from the spacer block.

	[image: image3.jpg]



	[image: image4.jpg]





2. Detach the spacer block from the body block.

	[image: image5.jpg]4

/

,,,m,./ : _'.r)/
// b&:&»ﬁ
%
..g..‘_‘.._\_\_;.. \v

\

-
» v.h

b 4
i
¥





	[image: image6.jpg]





3. Unscrew the 15mm outlet from the threaded hole in the right face of the body block.  Ensure the port and threads are free of foreign matter.  Replace the thin O-ring (p/n 0330214) and refit the 15mm outlet.  


	[image: image7.jpg]



	[image: image8.jpg]e
Ve¥ 50\

& viaw







1. Unscrew the hose from the right angle adapter and the right angle adapter from the flowmeter.  Remove the original PTFE tape from the right angle adapter, rewind with 2 complete layers of new PTFE tape and refit the adapter vertically into the flowmeter.

	[image: image9.jpg]



	[image: image10.jpg]





2. Unscrew the adjustable valve from the right hand threaded hole on the top face of the body block, and service the valve according to the Adjustable & Precision Valves Servicing Manual. Ensure the port and threads are free of foreign matter. Replace the thin O-ring (p/n 0330214) and refit the valve.

	[image: image11.jpg]



	[image: image12.jpg]e
Ve¥ 50\

& viaw







3. Unscrew the blanking bolts from the threaded left side hole and remaining threaded hole on the bottom face of the body block respectively.  Ensure the ports and threads are free of foreign matter.  Replace the thin O-rings (p/n 0330214) on the blanking bolts and refit into the body block.

	[image: image13.jpg]



	[image: image14.jpg]





4. Remove the precision valve.  Ensure the port and threads are free of foreign matter. If the valve has PTFE tape on the threads, remove it completely. Remove the thin O-ring, and replace with new O-ring (p/n 0330214). Note: do not replace the PTFE tape, as it is no longer required due a design change. Refit the precision valve into the centre bottom threaded hole.

	[image: image15.jpg]



	[image: image16.jpg]e
Ve¥ 50\

& viaw







5. Unscrew the Flowmeter, remove the top O-ring and flowmeter bolt from the left hand non-threaded hole in the body block.  Ensure the port and threads are free of foreign matter.  Replace the thin O-ring (p/n 0330214) on the lower part of the bolt. Replace the flowmeter bolt into the body block, fit a replacement thick O-ring (p/n 0330213) to the top of the bolt, and refit the flowmeter.

	[image: image17.jpg]



	[image: image18.jpg]



	[image: image19.jpg]e
Ve¥ 50\

& viaw







6. a). Unscrew the pressure gauge.  b). Ensure the port and threads are free of foreign matter.  c). Insert a new diamond copper washer into the single threaded hole on the label side of body block.  d). Screw in the pressure gauge and use the 14mm open ended spanner to tighten; the gauge should appear straight in comparison to the block and not able to be removed without the aid of tools. 

	[image: image20.jpg]



	[image: image21.jpg]e
Ve¥ 50\

& viaw







7. Attach the spacer block to the body block.

	[image: image22.jpg]4

/

,,,m,./ : _'.r)/
// b&:&»ﬁ
%
..g..‘_‘.._\_\_;.. \v

\

-
» v.h

b 4
i
¥





	[image: image23.jpg]





8.

Attach rail clamp to the spacer block.

	[image: image24.jpg]



	[image: image25.jpg]





9. Remove the original PTFE from the inlet hose threads and right angle adapter.  Line the thread of the inlet hose with two complete layers of new PTFE tape and refit the hose into the inlet hose right angle adapter using 12mm open-ended spanner ensuring it is tight. 

	[image: image26.jpg]



	[image: image27.jpg]





	Testing & Calibration: TT490-15.


Note: If adjustment/set up of adjustable or precision valves is required, valves must be removed and serviced according to Service Manual for adjustable or precision valves. Alternatively, instruments can be returned to Viamed for servicing, as specialist tooling is required to dismantle and reset.

1. Check the TT490-15 for damage.

2. Set the adjustable valve to minimum (fully counter clockwise) and the flowmeter to minimum (fully clockwise).

3. Connect the inlet hose to the oxygen supply at a pressure of 4 bar.

4. Connect the digital manometer (on 0-100.0 mbar range) with T piece adapter to the Tom Thumb outlet.

5. Set the flow meter to 15 Lpm.

6. Check all mechanical connections between parts for leaks; use Snoop fluid and check for bubbling.

7. Check for minimum pressure out: Cover the T piece adapter port. Check the TT pressure gauge is less than 8.0 cmH2O. Check that the digital manometer reading is less than 7.8 mbar and record. 

8. Set up the adjustable valve: Set the adjustable valve control to maximum. Cover the holes in the precision valve and the T piece adapter port. Alter the adjustable valve collar until a reading of 43 - 47 cmH2O is achieved on the TT pressure gauge. Ensure a reading of 42.2 – 46.1 mbar on the digital manometer and record. Slightly unscrew the adjustable valve and apply oxygen-compatible adhesive to the exposed threads. Retighten the adjustable valve with the adjustable spanner.

[image: image28.jpg]



9. Check for TT pressure gauge sticking: Release and cover the T piece adapter port a number of times, ensuring that TT pressure gauge returns to 43 - 47 cmH2O. Uncover the precision valve holes.

10. Set up the precision valve: Unscrew the precision valve adjustable screw until the thread is exposed 2 mm and apply oxygen-compatible adhesive to the exposed threads.  Screw into the precision valve body until flush. Cover the holes in the adjustable valve and the T piece adapter port.  Carefully adjust the precision valve screw in quarter turns to achieve a reading on the TT pressure gauge of 3.0 cmH2O greater than the reading recorded at (8) (Clockwise adjustment on precision valve screw to increase). Record the digital manometer reading.

[image: image29.jpg]



11. TT pressure gauge accuracy check: Cover the T piece adapter port and increase the flow of gas to achieve a displayed reading of 40 cmH2O on the TT pressure gauge. Ensure a reading of ± 2% of the digital manometer reading and record.

12. With the Flowmeter set at 15Lpm, gently shake the unit to check that the ball remains at the correct reading (not sticking), and then turn the Flowmeter and adjustable valve to minimum. Disconnect from oxygen supply and disconnect the digital manometer.

13. Clean the Tom Thumb with isopropyl alcohol.

14. Attach the TT490-15 Instructions for Use card (p/n 0390007).




TOM THUMB CALIBRATION / SERVICE SHEET
	Description.
	Tom Thumb.
	Serial No.
	

	Model.
	TT 490-15.
	Time & Date of Service / Calibration.
	


Record the manometer reading in millibars below, ensuring the TT490-15 meets the limits specified.

	Test Equipment
	Test
	Specification
	Reading
	P/F

	
	
	
	Calibration
	

	
	
	
	
	

	Snoop liquid.
	Check all ports / connections for leaks.
	No bubbling.
	
	

	
	
	
	
	

	Manometer
	Adjustable Valve:

@ 15 Lpm
	Minimum:

( 8 cmH2O.
	
	

	
	
	
	
	

	Manometer
	Adjustable Valve:

@ 15 Lpm.
	Tolerance:

( 43 & ( 47 cmH2O.
	
	

	
	
	
	
	

	Manometer
	Precision Valve:

@ 15 Lpm.
	Maximum:

+3.0 cmH2O over the adjustable valve setting.
	
	

	
	
	
	
	

	Manometer
	Pressure Gauge Test:

@ 40 cmH2O.  
	Tolerance: ± 2% 
	
	

	
	
	
	
	

	Visual check
	Flowmeter readings remain stable after shake test.
	
	
	

	
	
	
	
	

	Visual check.
	Gauge cannot be removed without tools.
	
	
	

	
	
	
	
	

	Visual check.


	Gauge appears straight.
	
	
	

	
	
	
	
	

	Visual check.
	All adjustable settings are set to a minimum.
	
	
	

	
	
	
	
	

	Visual check.
	Labels are attached.
	CE Label.
	
	

	
	
	Viamed Flowmeter Label.
	
	

	
	
	Serial Number Label.
	
	

	
	
	Tom Thumb Label
	
	


	Service.
	Parts Replaced.

	Changed O-Rings
	
	Diamond Copper Washer
	

	Re-greased Adjustable Valve Internal Parts
	
	Adjustable Valve
	

	Re-calibrated Tom Thumb
	
	Adjustable Valve Internal Parts
	

	
	
	Safety Valve
	

	
	
	Safety Valve Internal Parts
	

	
	
	Pressure gauge
	

	
	
	Pressure Gauge Face
	

	
	
	O-Rings
	

	
	
	Flow Gauge
	

	
	
	Flow Gauge Cover
	

	
	
	Outlet
	

	
	
	Inlet 
	

	
	
	Blanking Bolts
	

	
	
	Rail Clamp
	

	
	
	Main Body Block
	


Calibration: …………………………………
Signed:  …………………………………           















�    0086





Do not re-use the original “Diamond Copper Washer”.








© Copyright 2005. Viamed Ltd. All Rights Reserved
6

_1122380578

