	     [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley - West Yorkshire BD20 7DT - United Kingdom
Tel: +44 1535 634542                 Fax: +44 1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



Viamed VM-2500-S Series Capnograph
Serial number:

Firmware: 

RMA: 

Document No.: 

Number: 

Gas Accuracy Check
Service number: 4480520
	Test
	Range
	Displayed value EtCO2/FiCO2
	Result

	Calibration gas 

5% +/- 0.1vol% CO2, 20.9% O2, balance N2
	EtCO2: 4.7 - 5.3 vol%

FiCO2:  0.0 - 0.3vol%


	 /
	

	Breaths per minute
 
	10 (+/- 1bpm)
	
	

	ISA module leakage test
	CO2 drop 
<0.5vol% over 1 minute


	
	


	Visual inspection
	

	Power supply checks
	

	SpO2 functionality
	

	Alarms and data download functionality
	


Certificate number:

Signed:

Name:    

Position:  

Date:       

v1.0
[image: image1.png]