VANDAGRAPH
  COMPANY OPERATING PROCEDURE

Oxygen Sensor Manufacture

R-33 DE Assembly

VM3/COP/40.13

Date: 10-Mar-2006
Revision Date: 3-Jul-09
Issue 1

Only to be attempted by trained personel




Parts Required


Quantity
Description
Part No.


1
R33D1 Sensor



1
White Box 
9040021


1
White Box  Lid
9040002


1
JST SM02B-SRSS-TB socket  (515-1490)
9071014



Super glue



1
Packing Material
9040006


1
Round “Vandagraph” Label



2
Serial number Labels



1
Serial number Label for Outer bag



1
Polythene Bag (outer)



1
Installation instructions



Assembly: Set Temperature of Iron to 350 C Fine bit

[image: image1.jpg]


[image: image2.jpg]



1. [image: image3.jpg]


Remove the R33D1 Sensor from the Gas barrier bag, and carefully remove by unsoldering the Red  wire. Cut the balck wire as near to the PCB as possible. .
Do not apply too much heat

2. Hold the sensor in a Flow divertor which is held in a vice.

3. Solder the Red hole and centre holes

4. [image: image4.jpg]


Glue the JST socket onto the PCB Hold down until dry. 

5. The two terminals should line up with the outside of the solder points

6. [image: image5.jpg]


Ensure it is the correct way up
The rear terminals should point down towards the PCB And the narrowest part of the connector should be glued to the PCB








7. Solder the wires on the connector to the Red  and centre holes where the wires were removed and then to the socket 

8. Test the Sensor .The output should be between 23mV and 27mV. An Analox analyser can be used.

9. Remove the existing sensor label. 

10. Use printed labels with Citizen printer. R-33DE  : Serial number: Date code 

11. Carefully mount one label on the sensor and one label on the white container so that the edges of the labels are parallel to the edges of the container and Sensor.
Place packing material in the bottom and top of the white container and add the cap. Ensure it is sealed.

12. [image: image6.jpg]


[image: image7.jpg]


[image: image8.jpg]


[image: image9.jpg]


[image: image10.jpg]


Add the Vandagraph round label to the top of the cap. Add the instruction leaflet, and place the whole into the polythene bag, seal it and affix a Serial Number Label.

R-33DE JUly 2009.doc        7/3/2009         Page 2 of 2

