IEC 601-1 2nd EDITION 1988 MEDICAL
ELECTRICAL EQUIPMENT
PART 1: GENERAL REQUIREMENTS FOR SAFETY

	Clause No.
	Description, Requirements
	Compliance
OK N/A Fail
	Results, Observations, Notes/Comments

	6
	Identification, marking and documents
	/
	
	
	

	6.1
	Marking on the outside of equipment or equipment parts
	/
	
	
	

	6. le)
	Manufacturers or suppliers name and/or trademark
	/
	
	
	

	6.11)
	Model or type reference
	I’
	
	
	

	6. Ig)
	Rated Voltage range to which equipment may be connected
	
	I’
	
	

	
	Number of phases
	
	/
	
	

	
	Type of current
	
	I
	
	—____________________________________

	6.lh)
	Rated. supply frequency in Hz
	
	/
	
	

	6.lj)
	Power input (VA, W or A)
	
	I
	
	

	6.1k)
	Power output of auxiliary mains socket outlets
	
	/
	
	

	6.11)
	
Class II symbol
	
	1
	
	Not applicable to component.

	
	Symbol for degree of protection provided by enclosure e.g. IPX1, IPX4 or IPX7
	
	I
	
	Not applicable to component.

	
	Symbol for protection against electric shock


	
	I
	
	Not applicable to component.

	6.11m)
~-
	Mode of operation (if no marking, suitable for continuous operation)
	
	I
	
	Suitable for continuous operation.

	6. 1n)
	Types and ratings of external accessible fuses
	
	I
	
	No fuses.

	6.1p)
	Rated output voltage and current or power output and frequency
	
	~/
	
	Signal level outputs.

	6. lq)
	Physiological effects (symbols and warning statements)
	I
	
	
	Risk of burn or death if swallowed marking.

	6. 1r)
	Category AP/APG equipment (clause 38)



	
	
	
	Not applicable to component.

	6.1s)
	High Voltage symbol


	
	I
	
	Not a high voltage component.


	Clause No.
	Description, Requirements
	Compliance


OKN/A Fail
	Results, Observations, Notes/Comments

	6.1t)
	Special cooling requirements
	
	I
	
	

	6.1u)
	Limited mechanical stability
	
	I
	
	

	6.1v)
	Protective packaging markings
	
	1
	
	Info needed.

	6.1y)
	Potenti~1 equalization conductor terminal



	
	
	
	

	6.1z
	Functional Earth Terminal
	
	/
	
	

	
	Durability test of markings
	I
	
	
	

	6.2


6.2a)
	Marking on the inside of equipment or equipment parts


Is nominal supply voltage or voltage range of permanently installed equipment?
	
	I
	
	No markings within the equipment.

	6.2b)
	Is maximum power loading of heating elements or holders for heating lamps
	
	/
	
	No beating lamps.

	6.2c)
	High voltage parts
	
	if
	
	No high voltage parts.

	6.2d)
	Type of battery, polarity and mode of insertion
	
	1
	
	No batteries.

	6.2e)
	Fuses, type and rating or reference
	
	I
	
	No fuses.

	‘f)
---
	Protective earth terminal
	
	I
	
	No protective earth.

	ô.2g)
	
.j__
Functional earth terminal
	
	1
	
	No functional earth terminal.

	6.2h)
	Supply nmtral conductor in permanently
installed equipment

	
	1
	
	No neutral conductor.

	6.2j)
6.2k)

	Are markings on or near electrical connection points no: affIxed to parts which have to be removed to make the connection?
	
	I
	
	

	
	Are they visible after connection
	
	/
	
	

	
	Are the supply connections clearly marked adjacent to the terminals. For small equipmen description in accompanying document;?
	
	1
	
	No safety hazard involved if terminals are misconnected.

	Clause No.
	Description, Requirements
	Compliance
OKN/A Fail
	Results, Observations, Notes/Comments

	6.21)
	Is the equipment marked with the following statement (if at any point within a terminal box or wiring compartment intended for connection of the power supply conductors attains a temperature of more than 75°C during the normal temperature test):
‘For supply connections, use wiring materials suitable for at least ?
	
	I
	
	

	6.3
	Marking of controls and instruments
	
	I
	
	No mains switch on component.

	6.3a)
	Is the mains switch clearly identified? Markings of the positions of the mains switch
	
	
	
	

	6.3b)
	Is indication of different positions of control devices and different positions of switches on equipment (where patient hazard is possible) ?
	
	I
	
	No controls on component.

	6.3c)
	Indication of the direction of setting devices if safety hazard to the patient possible
	
	
	
	

	6.30
	Are controls and indicators with safety functions, e.g. alarms, identified?
	
	
	
	

	6.4
6.4a)
~__________
	Symbols
Compliance with appendix D, where applicable?
	
	
	
	

	6.4b)
	Do the symbols for controls and performance conform to IEC publication 878?
	
	
	
	

	6.5
6.Sa)
	Colors of insulation of conductors Is the protective earth conductor green/yellow throughout its length?
	
	
	
	

	6.Sb)
	Is insulation of internal conductors, which connect accessible metal parts to the protective earth terminal green/yellow?
	
	
	
	

	6.5c)
	Are only protective earth or potential equalization conductors green/yellow? (see 6.51) and 18 for exceptions)
	
	
	
	

	6.sd)
	Is the neutral conductor in the supply cord light blue?
	
	
	
	


	Clause No.
	Description, Requirements
	Compliance OKN/A Fail
	Results, Observations, Notes/Comments

	6.5e)
	Are the colors of conductors in the power supply cord in accordance with 1EC227 and
IEC 245?
	
	/
	
	No power supply conductor on component.

	6.51)
	Is the multi-conductor to Potential Equalization connection _ 0.1?
	
	I
	
	No bonding conductors.

	
	Is the end of the multi conductor colored or
marked green/yellow?
	
	I
	
	No bonding conductors.

	6.6

•


6.6a)
	Identification of medical gas cylinders and connections
Are these in compliance with ISO/R32?
	
	I
	
	Component not composed of gas cylinders.

	6.6b)
	Is the point of connection of gas cylinders so identified that errors are avoided when a replacement is made?
	
	I
	
	Component not composed of gas cylinders.

	6.7
6.7a)
	Indicator lights and push buttons Colored according to Table III?
	
	I
	
	No indicator lights.

	
	Red
	
	/
	
	No indicator lights.

	
	Yellow
	
	I
	
	No indicator lights.

	
	Green
	
	I
	
	No indicator lights.

	6.7b)
	Color red only used for emergency push buttons?
	
	I
	
	No indicator lights.

	6.8
~
	Accompanying documents
Are instructions for use delivered with the product?
	
	
	
	

	
	Is a technical description included with the equipment?
	
	
	
	

	
	Is a reference address provided
	
	
	
	

	
	Are warning statements and symbols explained in the accompanying documents?
	
	
	
	

	
	Is the language of the accompanying documents suitable for the destination country?
	
	
	
	



Date:
11/08/95

	Clause No.
	Description, Requirements
	Compliance OKN/A Fail
	Results, Observations, Notes/Comments

	6.8.2
6. 8.2a)
	Instructions for use
Do the instructions for use contain an explanation of the controls, displays and signals, sequence of operation, connection and disconnection of detachable parts and replacement of material(s) which is consumed during operation?
	
	
	I
	Info needed

	
	A description of the accessories?
	
	I
	
	No accessories.

	
	A description of device labelling?
	
	I
	
	

	
	Information on maintenance, cleaning and preventative inspections?
	
	I
	
	Does not apply to component evaluation.

	6.8.2c)
	Specified connection of equipment to
SIP/SOP?
	
	if
	
	Does not apply to component evaluation.

	6.8.2d)
	Parts in contact with patient:
	
	J
	
	

	
	Cleaning method
	
	,f
	
	

	
	disinfection method
	
	I
	
	

	
	sterilization method
	
	I
	
	

	6.8. 2e)
	Warning as to the necessity for periodical checking or replacement of additional or internal power source if not automatically maintained in a fully usable condition?
	
	I
	
	

	6.8.21)
	Warning to the user to remove batteries when not in use or if the equipment is not to be used for some time?.
	
	I
	
	No batteries.

	~.8.2g)
	Use and maintenance of rechargeable batteries?
	
	I
	
	No batteries.

	6.8.2h)
6.8.3
6.8.3a)
6.8.3b)
6.S.3c)
	Identification of appropriate power supplies or battery chargers?
	
	
	
	No power supply required.

	
	Technical description Performance data and characteristics details to enable safe operation?
	
	
	
	

	
	Particular measures or conditions for installation of the equipment?
	
	I
	
	

	
	Instructions for replacement of interchangeable and/or detachable parts (e.g. fuses in permanently installed equipment)?
	
	
	
	No replaceable parts.

	
	(Only for parts which are designated to be replaced). Has a statement of availability been made for:

Circuit diagrams?
	
	I
	
	No replaceable parts.



                                                   ~_•_.__       

	Clause No.
	Description, Requirements
	Compliance Results, Observations, Notes/Comments OKN/A Fail

	
	Description of function?


Service and calibration instructions
	
	I


I
	

	6.8.3d)
	Are any restricted environmental conditions for transport and storage specified?
	
	
	Verify

	7.
7. la)
	Power Input
Power consumed mainly by electric motor(s):
Allowed deviation:
+25% (rated input _ 100 W or 100 VA)
+15% (rated input > 100W or lOC) VA)
	
	I
	No power input.


Rated Value:


Measured value:

	7. ib)
	Power consumed mainly by other components:
Allowed deviation:
+15% (rated input _ 100 W or 100 VA)
+10% (rated input < 100W or 100 VA)
	
	I
	Rated Value:


Measured value:

	10.


10.1
	Environmental Conditions
Is equipment capable of, while packed for transport or storage, of being exposed for a period not exceeding 15 weeks to environmental conditions not outside of the following ranges:
	I
	
	

	10.1.a)
	Ambient temperature range of -40°C to
+70°C
	I
	
	

	10.1.b)
	Relative humidity range of 10% to 100%, including condensation.
	1
	
	

	10.1.c)
	Atmospheric pressure range of 500hPa to lO6OhPa.
	I
	
	

	.2.1
~
	Environmental requirements for test conditions.
	I
	
	

	10.2.2
	Power Supply
Power supply voltages and frequency.
	
	I
	No power supply required.

	14.
	Requirements related to classification
	
	I
	Not class I equipment.

	14.1
	Class I equipment
	
	
	

	14. ib)
	If the mains part of equipment specified for an external d.c. power source is isolated from accessible conductive parts by basic insulation only, is a separate protective earth conductor provided?
	
	
	

	14.2 l4.2a)
	Class Ii equipment

Insulation enclosed?
	I
	
	Component is SELV.

	
	Metal enclosed?


Combination of above?
	
	
	

	-
	
	
	
	

	Clause
	Description, Requirements
	Compliance
	Results, Observations, Notes/Comments

	No.
	
	OKl’~/A Fail
	

	14.2b)
	Is there a device for changing over from class to class II protection?
	
	I
	
	No device for converting from class Ilto Class I.

	
	If yes, are all following requirements fulfilled:
	
	
	
	

	
	Clear indication of the selected class?
	
	
	
	

	
	Use of a tool is necessary for change over


In class II position, is PB connection interrupted or changed into a functional earth connection?
	
	/


I
	
	

	.
	Compliance of the equipment with the requirements of the selected class given at any time?
	
	I
	
	

	14.4
	Class I and 11 equipment
	I
	
	
	SELV.

	14.4a)
	Is the equipment provided with an additional protection according to the requirements of class I or class II equipment
	
	
	
	

	14.4b)
	Equipment for connection to external d.c. power supplies:
Does no hazard arise if the supply polarity is reversed?
	
	I
	
	.

	14.5
	Internally powered equipment
	
	I
	
	No means for connection to a supply mains.

	14.5a)
	Is equipment with internal power source with means of connection to a supply mains dual classified?
	
	
	
	

	~i4.5b)
	Does the equipment intended for connection to a supply mains comply with the requirements for class I or II equipment while so connected?
	
	/
	
	No means for connection to a supply mains.

	14.6
	Type B, BF and C’F equipment
	
	I
	
	Not intended for direct cardiac application.

	14.6c)
	Is the equipment intended for direct cardiac application type CF?
	
	
	
	

	14.ód)
	Does equipment with applied parts of type CF and type B or BF comply with the
	
	/
	
	No applied parts.

	
	requirements of 6.11?
	
	
	
	

	IS.
	Limitation of voltage and/or energy
	
	I
	
	Component not intended to be connected to

	l5.b)
	Are residual voltages between supply plug pins and between supply pin and enclosure after is _ 60 V?
	
	
	
	the supply mains.

	
	
	
	
	
	

	Clause No.
	Description, Requirements
	Compliance
OKN/A Fail
	Results, Observations, Notes/Comments


Retention of charge in capacitor. Residual voltage: _ 60 V?

I

	Residual energy: _ 2 mJ?
	
	I
	Component not intended to be connected to the supply mains.

	Endosures and Protective Covers


Protection against accidental contact with live parts and parts which can become live in the event of failure of basic insulation?
	I
	
	Does not apply to component. Operates at
SELV.

	Are openings in top covers
	of enclosure
	so
	Does not apply to component. Operates at

	positioned
	or dimensioned
	that accessibility to
	SELV.

	live parts
	is prevented (test
	rod)?
	

	Do conductive parts of actuating mechanisms of electrical controls which are accessible, after the removal of handles, knobs and levers, have either a resistance _ 0.2 to protective earth or does separation from live parts comply with clause 17g?
	I
	No accessible conductive parts.

	Are internal parts of the equipment with a circuit voltage exceeding 25 VAC or 60 VDC, which cannot be disconnected from the supply, protected against contact even after opening of the enclosure?
	I
	No circuits over 25 VAC or 60 VDC.


OR

	Are protective enclosures removable only with the aid of a tool?
	I
	Protective covers are not removable.

	Is there an automatic device which disconnects
/ accessible live parts from the supply when the
	No live parts accessible.

	enclosure
	is opened?.
	
	

	Are live parts inaccessible to the test rod through openings for adjustment of pre-set controls?
	
	
	
	No openings.

	Separation


Separation method of the applied part from live pans:


Are allowable leakage currents exceeded?
	
	I
	
	Separated from mains part or other live parts (please specify):


No applied parts. Components are not hand held.

	Basic insulation


Applied part protectively earthed?
	
	I
	
	No applied parts. Components are not hand held.

	By protectively earthed conductive part (e.g. screen)?
	
	/
	
	No applied parts. Components are not hand held.


	Clause No.
	Description, Requirements
	Compliance OKN/A Fail
	Results, Observations, Notes/Comments

	17.a)3)
	By separate earthed intermediate circuit?
	
	I
	
	No applied parts.

	17.a)4)
	By double or reinforced insulation?
	
	I
	
	No applied parts.

	17.a)5)
	By protective impedances limiting current to applied part?
	
	I
	
	No applied parts.

	17.a)
	Other method; e.g. specified in particular standard
	
	I
	
	Dei~i~
FauI~
Cc~ditic*i
	Strth
Leakage
Cuncn~
	Enc1~urc
Le*ka~c
Cur,~
	Pat~t
Lca~cigc
Cwyet,t

	
	Additional leakage current tests in single fault conditions (see clause 19.4)
	
	
	
	
	
	-
	

	17.c)
	Is there no connection between applied parts and accessible conductive parts, which are not protectively earthed?
	
	I
	
	No applied parts.

	17.d)
	Supplementary insulation between hand-held flexible shafts and motor parts (class 1)
	
	I
	
	Not a hand held component. No motors.

	
	Rated motor voltage:
	
	I
	
	V

	
	Test voltage:
	
	I
	
	V

	
	Adequate mechanical strength:
	
	I
	
	
	
	
	

	
	Air clearances/creepage distances:
	
	I
	
	minI
mm

	17.g)
	Separation methods of accessible parts (other than applied parts) from live parts in normal and single fault condition:
(Allowable leakage currents are not exceeded)
	I
	
	
	Component operates as a SELV.

	
	I.
Basic insulation - accessible part protectively earthed?
	
	I
	
	Component operates as a SELV.

	
	2.
By protectively earthed conductive part (e.g. a screen)
	
	I
	
	Component operates as a SELV.

	
	3.
By separate earthed intermediate circuit?
	
	I
	
	Component operates as a SELV.

	
	4.
By double or reinforced insulation?
	
	I
	
	Component operates as a SELV.

	
	5.
By protective impedances limiting current to the accessible part?
	
	I
	
	Component operates as a SELV.

	18.
	Protective earthing, functional earthing and potential equalization
	
	I
	
	Not a class I type of equipment.

	18.a)
	Are accessible parts of class I equipment

	
	which are
	separated from live parts by bask

	
	insulation,
	connected by a sufficiently low

	
	impedance
	to the protective earth conductor?

	Clause No.
	Description, Requirements
	Compliance
OKN/A Fail
	Results, Observations, Notes/Comments

	18.b)
	Is the protective earth terminal suitable for connection to the protective earth conductor of the power system?
	
	I
	
	No protective earth terminal.

	18.e)
~
	If the equipment is provided with means for the connection of a potential equalization conductor, does this connection comply with the following requirements:


The connection is readily accessible?
	
	/
	
	No potential equalization conductors.

	
	Accidental disconnection is prevented in normal use?
	
	I
	
	
	
	

	
	The conductor can be detached
without the use of a tool?
	
	I
	
	
	
	

	
	Power supply cord does not incorporate a potential equalization conductor?
	
	I
	
	
	
	

	
	Connection is marked with symbol 9 of table DI?
	
	I
	
	
	
	

	18.0


~,
	Impedance of protective earthing system Is the impedance _ 0.1 between the PE terminal and any other earthed part where a flexible supply cord is not used or where a detachable supply cord is used OR


is the impedance _ 0.2 between the protective earth terminal in the mains plug and any other earthed part where a non-detachable supply cord is used?
	
	I
	
	No protective earth terminal.

	
	Measurement between any accessible part and:
	
	
	
	Measurement
	Requirement

	
	Protective Earth Terminal
	
	I
	
	No protective earth.
	_ 0.1

	
	—
Appliance inlet
	
	I
	
	No appliance inlet.
	_ 0.1

	
	Mains plug
	
	I
	
	No mains plug.
	_ 0.2

	18.g)
	If the impedance of protective earth connections other than in 18.1) exceeds 0.1, does the continuous fault current to an accessible part not exceed the allowable value of the enclosure leakage current in single fault condition?
	—
	I
	
	No protective earth.

	18.k)
	Are functional earth terminals not used to provide functional earthing?
	
	I
	
	No protective earth.


	Clause No.
	Description, Requirements
	Compliance OKN/A Fail
	Results, Observations, Notes/Comments

	18.1)
-~
	For class II equipment with isolated internal screens and with power supply cord (three conductors):
	
	
	I
	
	No power supply cord.

	
	Is the third conductor only used as functional earth of these screens, and is it colored green/yellow?
	
	
	I
	
	No power supply cord.

	
	Is the insulation of these screens, and all internal wiring connected to them, double or reinforced insulated?
	
	
	I
	
	No power supply cord.

	
	Is the marking of the functional earth terminals distinguished from protective earth terminal and is it noted in the accompanying documents?
	
	
	/
	
	No power supply cord.

	19.
	Earth Continuity, Earth Leakage Currents, and Insulation Resistance
	
	
	I
	
	Result
	Limit(s)

	
	Protective Earth Continuity
	
	I
	
	No protective earth.
	0.100

	INSULATION POTENTIAL 500 Volts dc
	
	
	
	
	
	

	
	Insulation Resistance Mains Fl to Case
	
	
	I
	
	
	
	>50M

	
	Insulation Resistance Mains F2 to Case
	
	
	I
	
	
	
	> 50M

	
	Insulation between mains & Applied Parts
	
	
	I
	
	
	
	> SOM

	
	insulation between Applied Parts and Case
	
	
	I
	
	
	
	>SOM

	Earth Leakage Current

	
	Earth leakage Current SFC; Open Supply
	
	I
	
	
	A
	1000 A

	
	Earth leakage Current Normal Condition
	
	I
	
	
	A
	500 A

	
	Earth leakage Current SFC; Open Supply Mains reversed
	
	I
	
	
	A
	1000 A

	
	Earth leakage Current Normal; Mains reversed
	
	I
	
	
	A
	500 A


___ ENCLOSURE LEAKAGE CURRENT

Enclosure Leakage Current SFC; Open Supply

I

A

	Enclosure Leakage Current Normal Condition
	
	/
	
	A
	100 A

	Enclosure Leakage Current SFC; Open Circuit Earth
	
	I
	
	A
	500 A


Enclosure Leakage Current SFC; Open

Supply, Mains Reversed

500A

I
A

500A

Earth, Mains Rev

	Clause No.
	Description, Requirements
	Compliance OKN/A Fail
	Results, Observations, Notes/Comments

	
	
	Enclosure Leakage Current Normal Condition; Mains Reversed
	
	
	I
	
	
	A
	100 A

	
	
	Enclosure Leakage Current SFC; Open Circuit Earth Mains Rev
	
	
	I
	
	
	A
	500 A

	PATIENT LEAKAGE CURRENT
	

	
	Patient Leakage Current SFC; Open Supply
	
	
	I
	
	
	A
	

	
	Patient Leakage Current Normal Condition
	
	
	/
	
	
	A
	500 A

	
	Patient Leakage Current SFC; Open Circuit Earth
	
	
	I
	
	
	A
	100 A
500 A

	
	Patient Leakage Current SFC; Open Supply, Mains Reversed
	
	
	I
	
	
	A
	500 A

	
	Patient Leakage Current Normal Condition, Mains Reversed
	
	
	/
	
	
	A
	100 A

	
	Patient Leakage Current SFC; Open Circuit Earth, Mains Rev
	
	
	I
	
	
	A
	500 A

	
	Patient Leakage Current; Mains on Applied Part
	
	
	I
	
	
	A
	5000 A

	
	Patient Leakage Current; Mains on Applied Part Mains Rev
	
	
	I
	
	
	A
	5000 A

	PATIENT AUXILIARY CURRENT
	
	
	
	
	
	
	
	
	


Patient Auxiliary Current SFC; Open Supply
I

A

500A

	Patient Auxiliary Current Normal Condition
	I
	
	

	Patient Auxiliary Current SFC; Open Circuit Earth
	I
	A
	10 A
500 A

	Patient Auxiliary Current SFC; Open Supply, Mains Reversed
	I
	A
	500 A

	Patient Auxiliary Current Normal Condition, Mains Reversed
	I
	A
	10 A


A
500A

Report No. 556395 Page 13 of 44 Date: 11/08/95

	Clause
No.
—
	Description, Requirements
	Compliance
OKN/A Fail
	Results, Observations, Notes/Comments

	20
	Dielectric strength
(prior to humidity preconditioning)
	
	I
	
	Exception applies:
insulation is not a safety function.
	

	Equipment
	Insulation under test
	Reference Voltage
[V]
	
	
	
	Test voltage
[VI
	Observations

	All types
	A-al ~b.i.)
	
	
	I
	
	
	

	
	A-a2 (d.i. or r.i.)
	
	
	/
	
	
	

	
	A-b (b.i.)
	
	
	I
	
	
	

	
	A-c(s.i.)
	
	
	I
	
	
	

	
	A-e (d.i. or r.i.)
	
	
	I
	
	
	

	
	A-f (b.i.)
	
	
	/
	
	
	

	
	A-g (b.i. or r.i.)
	
	
	/
	
	
	

	
	A-j (s.i.)
	
	
	I
	
	
	

	
	A-k (d.i. or r.i.)
	
	
	I
	
	
	

	Types with
applied part
	B-a (d.i or r.i.)
	
	
	I
	
	
	

	
	B-b see particular standard
	
	
	I
	
	
	

	
	B-c (s.i.)
	
	
	I
	
	
	

	
	B-d (b.i.)
	
	
	I
	
	
	

	
	b-e (d.i. or ri.)
	
	
	I
	
	
	

	
	Overall compliance with clause 20?

I
b.i.
= basic insulation d.i. = double insulation r.i. = reinforced insulation s.i. = supplementary
insulation

	Note
	

	Clause
No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	20
	Dielectric strength
(prior to humidity preconditioning)
	
	I
	
	
	

	Equipment
	insulation under test
	Reference Voltage
I[VI
	
	
	
	Test voltage
[VI
	Observations

	All types
	A-al (b.i.)
	
	
	/
	
	
	

	
	A-a2 (d.i. or r.i.)
	
	
	/
	
	
	

	
	
	
	
	
	
	
	

	
	A-b(b.i.)
	
	
	I
	
	
	

	
	A-c(s.i.)
	
	
	/
	
	
	

	
	A-e (d.i. or r.i.)
	
	
	/
	
	
	

	
	A-f(b.i.)
	
	.
	I
	
	
	

	
	A-g (b.i. or r.i.)
	
	
	I
	
	
	

	
	A-j (s.i.)
	
	
	I
	
	
	

	
	A-k (d.i. or r.i.)
	
	
	I
	
	
	

	Types
with
applied part
	B-a (d.i or r.i.)
	
	
	I


I
	
	.
	.

	
	B-b see particular standard
	
	
	
	
	
	

	
	B-c (s.i.)
	
	
	I
	
	
	

	
	B-d (b.i.)
	
	
	I
	
	
	

	
	b-e (d.i. or r.i.)
	
	
	I
	
	
	

	
	Overall compliance with clause 20?
	
	I
	
	
	

	Note



	—


b.i. = basic insulation d.i. = double insulation r.i. = reinforced insulation s.i. = supplementary insulation

	
	
	
	
	
	
	
	

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comn~ents

	21.
	Mechanical Strength


Are enclosures, including any access covers forming part of them, of sufficient strength and rigidity?
	I
	
	
	Polymeric enclosure is suitable for intended application.

	21.3
	Is there no damage to parts of patient support and/or immobilization system after loading test (1.35 Kn and 2.7 Kn)?
	
	I
	
	No patient immobilization components.

	21.5
	Are hand held equipment or equipment parts safe after drop test (drop height 1 m)?
	
	I
	
	Not hand held during normal use.

	21.6
	Is portable and mobile equipment able to withstand rough handling?
	I
	
	
	Components are able to withstand a drop from Scm onto a hardwood floor.

	22.


22.2a)
	Moving Parts


Are moving parts provided with guards which form an integral part of the equipment (transportable equipment)?
	
	I
	
	No moving parts.

	22.2b)
	Are moving parts provided with similar guards as above unless it is evident that equivalent protection will be separately provided during installation (stationary equipment)?
	
	I
	
	No moving parts.

	22.3
	Are cords, chains, belts, etc. so confined that they cannot run off or jump out of their guiding devices?
	
	I
	
	No moving parts.

	
	Are guides or other safeguards removable only with the use of a tool?
	
	I
	
	No moving parts.

	22.4
	Are dangerous movements of equipment parts, which may cause physical injury to the patient only possible under the control of the operator?
	
	I
	
	No moving parts.

	22.6
	Are parts of equipment subject to mechanical wear accessible for inspection?
	
	I
	
	No moving parts.

	22.7
	Is there a means for emergency switching off a relevant part, to remove an unexpected safety hazard caused by an electrically produced mechanical movement?
	
	I
	
	No moving parts.

	
	Is the means of emergency switching readily identifiable and accessible and does it not introduce a further safety hazard?
	
	I


/
	
	No moving parts.

	
	Are the devices for emergency stopping able to break the full load current?
	
	
	
	No moving parts.


	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	22.7cont
	Does means for stopping of movements operate as the result of one single action?
	
	I
	
	No moving parts.

	23.
	Surfaces, Corners and Edges


Is the equipment free from sharp edges, corners and burrs?
	I
	
	
	

	24.


24.1
	Stability in Normal Use


Does the equipment not overbalance when tilted through an angle of 100 in normal use?
	I
	
	
	

	24.3
	if the equipment overbalances when tilted through 100, does it meet the following requirements:

does the equipment not overbalance when tilted through 50 in any position of normal use?
	
	I
	
	

	
	does the equipment carry a warning notice for transport?
	
	I
	
	

	
	Does the equipment not overbalance when tilted through 100 in the position specified for transport?
	
	I
	
	

	24.6


24.6a)
	Grips and orher handling devices


Equipment or its parts with a mass of more than 20 Kg:

is it provided with suitable handling devices?
	
	I
	
	No grips or handling devices.

	
	Is it provided with handling
instructions for lifting and
assembling?
	
	I
	
	No grips or handling devices.

	24.6b)
25.
25.1
25.2


~
	Portable equipment with a mass of more than 20 Kg:

does it have carrying handles, which are suitably place so that the equipment can be carried by 2 or more persons?
	
	I
	
	No more than 20 Kg.

	
	Expelled Parts


Are protective means provided where expelled parts of the equipment could constitute a safety hazard?
	
	I
	
	No expelled parts.

	
	Is a graphical display vacuum tube with a diameter of 16 cm or larger intrinsically safe with respect to effects of implosion and mechanical impact? OR
	
	I
	
	No graphical displays.


	lause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	
	Does the enclosure provide an adequate protection against implosion?
	
	/
	
	No implosion hazard.

	28.


28.3
	Suspended Masses


Suspension systems with safety devices

Do suspension systems with safety devices with adequate safety factors protect the user or patient from hazards? If so, is it obvious to the user when the safety device comes into operation?
	
	/
	
	No suspended masses.

	28.4
.
	Suspension systems of material without safety devices
1.
Does the total load not exceed the safe working load?
	
	I
	
	No suspended masses.

~

	
	2.
Where supporting characteristics are not impaired by wear, corrosion, material fatigue or ageing,, is the safety factor of supporting parts not less than 4?
	
	I
	
	No suspended masses.

	
	3.
Where impairment by wear, corrosion, material fatigue or ageing is expected, do the relevant supporting parts have a safety factor not less than 8?
	
	I
	
	No suspended masses.

	
	4.
Where metal with a specific elongation at break of less than 5% is used, do supporting parts have safety factors as in 2 and 3. above multiplied by 1.5?
	
	I
	
	No suspended masses.

	
	5.
Are sheaves, sprockets, bandwheels and guides so designed that the safety factors of these subclauses are maintained for a specific minimum life till replacement of the ropes, chains and bands?
	
	I
	
	No suspended masses.

	29


29.2
	X-radiation


Have precautions been taken to proI~ect the user/patient from harmful extraneous radiation from the equipment?
	—
	I
	
	No X-radiation

	
	
	
	
	
	


II


38.
Marking, Accompanying Documents

38.2

Is APG equipment marked prominently with a green band imprinted with the characters “APG~ (see Appendix DII)?

I

If this marking is impossible, is the relevant

information given in the instructions for use?

I

	38.4
	Is AP equipment marked prominently with a green circle imprinted with the characters “AP” (see appendix DII)?
	I

	
	If this marking is impossible, is the relevant information given in the instructions for use?
	I

	38.5
	Is the above marking (38.2 and 38.4) present on the major part of the equipment?
	

	38.6


~
	Do the accompanying documents contain an indication to distinguish the parts of the equipment which are AP and APG?
	I

	601-1 2nd EDITION 1988 MEDICAL ELECTRICAL EQUIPMENT

1: GENERAL REQUIREMENTS FOR SAFETY
	
	
	

	Clause
Description, Requirements
Compliance

0KM/A Fail
	Results, Observations, Notes/Comments

	cont Ionizing radiation emitted by the equipment

(other than X-ray equipment) utilizing vacuum tubes excited by voltages exceeding 5 KV:

Allowed value:

130 nC/Kg = 0.5 mR in 1 hour at a distance of 5 cm from the equipment.
	I
	
	No X-radiation.

	Alpha, beta, gamma, neutron, radiation and
other particle radiation (under consideration)
	I
	
	No X-radiation.

	Microwave radiation (under consideration)
	I
	
	

	Light radiation (including lasers)
	
	
	

	—
Infra red radiation (under consideration)
	I
	
	

	Ultraviolet radiation (under consideration)
	I
	
	

	Acoustic energy (including ultrasonics)
(under consideration)
	I
	
	

	Electromagnetic compatibility
JEC 62A(CO)41
	I
	
	

	Locations and basic requirements


Flammable anaesthetic mixture with air
	I
	
	No flammable anaesthetic mixtures with air.

	Flammable anaesthetic mixture with oxygen or
nitrous oxide
	I
	
	Flammable anaesthetic mixture with oxygen or nitrous oxide

	Are equipment or parts of the equipment, which are used in a location defined in 37.5, AP or APG equipment?
	I
	
	

	Clause No.
	Description, Requirements
	Compliance Results, Observations, Notes/Comments 0KM/A Fail

	38.7
	Where only parts of the equipment are AP or
	
	I
	Not AP or APG equipment.

	
	APG, are the relevant parts clearly marked?
	
	
	

	39.
	Common requirements for Category AP and Category APG equipment.
	
	I
	Not AP or APG equipment.

	39. la)
	Are creepage distances and clearances between the connection points of power supply cords as required (57.10 table XVI)?
	.
	
	

	39. lb)
	Are connections prevented from accidental disconnection or only removable with the use of a tool (except circuits complying with 40.3 and 41.3)?
	
	I
	Not AP or APG equipment.

	39. ic)
	Is category AP and APG equipment not provided with a detachable supply cord? (except circuits complying with 40.3 and 41.3)
	
	
	Not AP or APO equipment.

	39.2
	Construction details
	
	I
	Not AP or APG equipment.

	39.2a)
	Are covers for protection against penetration of gasses only removable with the use of a tool?
	
	
	

	39.2b)
	Do openings in covers meet the following requirements:
1.
Openings for control elements in the top cover are fully covered by the control knobs?
	
	I
	Not AP or APG equipment.

	
	2.
In the sides they do not allow penetration from a > 4mm test rod?


3.
In the base plate they do not allow penetration from a > 12 mm test rod?
	
	I


I
	Not AP or APG equipment. Not A? or APG equipment.

	39.2c)
	Are inadmissible temperatures or safety hazards due to a conductor short circuit to a conductive part, which contains flammable anaesthetic gas, prevented?
	
	I
	Not AP or APG equipment.

	39.3
	Prevention of electrostatic charges
	
	I
	Not AP or APG equipment.

	39.3a)
	Have measures been taken to prevent electrostatic charges by providing leakage paths to earth or by the use of anti static materials?
	
	
	

	39.3b)
	Do the electrical resistance limits of
	
	I
	Not A? or APG equipment.

	
	anaesthetic tubes, mattresses and pads, castor tires and other anti static materials comply
	
	
	

	
	with ISO standard 2882?
	
	
	

	
	
	
	
	

	Clause No.
	Description, Requirements
Compliance
0KM/A Fail
	Results, Observations, Notes/Comments

	39.4
	Corona
Are parts of equipment operating at more than 2000 V a.c. or 2400 V d.c., which are not included in an enclosure, so designed that corona cannot be produced?
	I
	
	Not A? or APG equipment.

	40.


40.1
	Requirements and Tests for Category A? Equipment, parts and components thereof Does equipment or equipment parts not ignite flammable anaesthetic mixtures with air in normal use and normal condition?
	I
	
	Does not apply to component evaluation. Must be investigated in final assembly.

	40.2
	Temperature limits

Does equipment, producing no sparks, remain under the following surface temperatures, where vertical air circulation is:


1.
restricted - 150°C?
	I
	
	Does not apply to component evaluation. Must be investigated in fmal assembly.

	
	2. unrestricted - 200°C?
	/
	
	Does not apply to component evaluation. Must be investigated in final assembly.

	40.3
	Low energy circuits


Does equipment which may produce sparks in normal use comply with:


1.
temperature requirements and
	I
	-
	Component does not produce sparks in normal use.

	
	2.
requirements of U, I, L and C?
	I
	
	

	40.4


4O.4a)
	External ventilation with internal overpressure


Are flammable anaesthetic mixtures with air removed by ventilation before the equipment is energized?
	I
	
	Does not apply to component evaluation. Must be investigated in final assembly.

	
	Is the penetration of such mixtures prevented by maintenance of overpressure within the equipment, by means of air not containing flammable gases?
	I
	
	Does not apply to component evaluation. Must be investigated in final assembly.

	40.4b)
	Is the ventilating gas in the enclosure at an overpressure of at least 0.75 hPa?
	I
	
	

	
	Is the energizing of equipment only possible if the required minimum overpressure has been present, so that the displaced volume is at least 5 times the volume of the enclosure?
	
	
	

	4O.4c)


j~
	Is the equipment provided with a pressure sensor to detect and to de~energize ignition sources when the overpressure drops below 0.5 hPa?
	I
	
	

	Clause No.
	Description, Requirements
Compliance
Results, Observations, Notes/Comments
0KM/A Fail

	42.


42.1
	Excessive Temperatures
.


Do equipment parts having a safety function and their environment not attain temperatures exceeding the values in Table Xa during normal use/normal condition over the range of ambient temperatures specified in 10.2.1?
	/
	Does not apply to component evaluation. Must be investigated in final assembly.

	
	Measuring point
	I
	Measured
AllowedT[°C]
T[°C]

	42.2
	Do equipment parts and their environment not attain temperatures exceeding the values given in Table Xb during normal use/normal condition at an ambient temperature of 25°C?
	I
	
	

	
	Measuring point
	
	Measured
T[°C]
	Allowed
T[°C)

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	42.2
cont
	Determination of the temperature rise of copper motor windings by the resistance method:


~t= (R2-R1/R1)x(234.5 + ti) + (t2-tl)
	
	I
	
	Insulation class:
Room temperature: ti =
Room temperature: t2 =
t=

	
	Overall compliance with clause 42?
	
	
	
	

	
	
Motor Details:
Manufacturer:
-
Part or Model No:
Rated Voltage:
Rated Current:
No. of Phases:
RPM Specified:
HP:


Motor Capacitor Ratings (if applicable)
Manufacturer:
Part or Model No:
Rated Voltage:
Capacitance:


Construction Details:

	
	TEMPERATURE RISE OF MOTOR WINDINGS BY CHANGE OF RESISTANCE METHOD

	
	R1WI=
1=
A
R2WI=
~T=____ Or

	
	R1W2=
1=’
A
R2W2=
~T=
OC,

	
	Ri W3 =
I =
A
R2 W3=
=~T
—____

	
	Ri W4 =
I
A
R2 W4=
=*T = —____

	
	R1W5=
1=
A
R2WS=
=~T=____

	
	R1W6=
1=
A
R2W6=
*T=____

	
	R1W7=
1=
A
R2W7=
~*T=

	
	R1W8=
1=
A
R2W8=
=~T=
°C



42.3
Surface temperatures

Are surface temperatures of applied parts, which are not intended to supply heat, not higher than 41°C?


I
Does not apply to component evaluation.



Must be investigated in final assembly.

	43.


43.1
	Fire Prevention


Does equipment have the strength and rigidity necessary to avoid fire hazards (see clause
21)?
	I
	
	
	

	44.


44.2
	Overflow, Spillage, Leakage, Humidity, Ingress of Liquids, Cleaning, Sterilization and Disinfection


Overflow


If equipment contains a liquid reservoir:
a)
is the equipment electrically safe in the event of 15% overfill?
	
	I
	
	No liquid reservoirs.

	
	b)
as for (a) and tilted through 15° for transportable equipment?
	
	I
	
	No liquid reservoirs.

	44.3
	Spillage


Does spillage not wet parts, which may cause a safety hazard?
Test:
A quantity of 200 ml of water is poured steadily on an arbitrary point on the top surface of the equipment for 15 seconds.
	
	I
	
	No liquids.

	44.4
44.5
	Leakage


Does liquid, which escapes in single fault condition, not wet parts, which may cause a safety hazard?


Test:
By means of a pipette, drops of water are applied to couplings, which might rupture.
	
	I
	No liquids.


— Not adversely affected by humidity.

	
	Humidjr~
Is the equipment sufficiently proofed against the effects of humidity:
Compliance is checked by preconditioning treatment
	—
	I
	


Compliance

0KM/A Fail

42.5

I

Results, Observations, Notes/Comments

Does not apply to component evaluation. Must be investigated in final assembly.

Guards

Are guards provided where there are hot accessible surfaces?

Are guards removable only with a tool
I

Does not apply to component evaluation.


Ingress of liquids

Is protection against harmful ingress of liquids

in accordance with IEC publication 529?
Drip-proof equipment test?
Splash-proof equipment test?

_________ Watertight equipment test?


44.7
Cleaning, Sterilization and disinfection



Sterilization or disinfection method specified



by the manufacturer:

b)
does the equipment satis& the

requirements of the dielectric strength

___________
test?


45.
Pressure vessels subject to pressure


45.2
Is the equipment covered under the inspection



procedures of a national regulation?

If yes, test 45.3 are not required.

Does the maximum pressure to which a part can be subjected in normal and abnormal operation not exceed the maximum permissible working pressure for the part? The used maximum pressure is the highest of

_________ the following:

4S.3a)
Rated maximum supply pressure from the


external source

45.3b)
The pressure setting of a pressure-relief

________
device provided as part of the assembly

Clause

No.

44.6

Description, Requirements

I

Compliance Results, Observations, Notes/Comments 0KM/A Fail

Does not apply to component evaluation.

I

I

I

I

	Test 20 cycles?
	I

	No sterilization or disinfection method is specified:
	/

	Test conditions: Steam 134 ± 4°C, 2 bar, 20 cycles, 20 mm duration each
	I

	After the above specified methods:


a)
are there any signs of deterioration?
	I


I

45.3

Not a pressure vessel.

I

I

45. 3c)

/

I

The maximum pressure which can be developed by an air compressor that is part of the assembly, unless the pressure is limited by a pressure relief device

I

No air compressors.

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	45.7


45.7a)
	Are pressure relief devices (where fitted):


Located as close to the pressure vessel or parts of the system they are intended to protect?
	
	I
	
	No pressure vessels.

	45.7b)
	Accessible for repairs, inspection and maintenance?
	
	I
	
	

	45. 7c)
	Cannot be adjusted or rendered inoperative without the use of a tool?
	
	I
	
	-

	45.7d)
	Is the discharge opening so located and directed that the released material is not directed towards any person?
	
	I
	
	

	45.7e)
	Is the discharge opening so located and directed that operation of the device will not deposit material on parts causing possible safety hazards?
	
	I
	
	

	45.70
	Is there an adequate discharge capacity?
	
	I
	
	

	45.7g)
	Is there no shut-off valve between a pressure relief and the part that it is intended to protect?
	
	I
	
	

	45.7h)
	Minimum number of cycles of operation of safety device: 100,000
	
	I
	
	

	49.


49.1
	Interruption of the power supply


If thermal cut—outs and overcurrent releases with automatic resetting are used, do they not cause a safety hazard by such resetting?
	
	I
	
	Does not apply to component evaluation.

	49.2
	Does interruption and restoration of the power supply cause no hazards?
	
	I
	
	

	49.3
	Are means provided for the removal of mechanical constraints on patients in the event of a supply mains failure?
	
	I
	
	Not a patient applied device.

	51.


51.1
	Protection against hazardous output Intentional esceeding of safety limits


Are means provided which prevent or indicate to the operator that the selected setting is in excess of a safety limit?
	
	I
	
	Does not apply to component.

	51.2


~
	Indication of parameters relevant to safety Does equipment delivering energy or substances to patients indicate a hazardous output?
	
	I
	
	Component does not deliver substances or energy to patient.

	Clause
Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations; Notes/Comments

	Accidental selection of excessive output values


In therapeutic equipment providing low and high outputs, are means providing to avoid the accidental selection of high output?
	
	I
	
	Not a therapeutic piece of equipment.

	Abnormal Operation and Fault Conditions


Is the equipment so designed and manufactured that in single fault condition, no safety hazards exist? (see clause 3.1 and clause 13)
	I
	
	
	

	
The following safety hazards shall be taken into consideration:


1
Emission of flames, molten metal, poisonous or ignitable gas in hazardous quantities.
	
	
	
	

	deformation of the enclosure
	
	
	
	

	temperatures exceeding the maximum values shown in table XII
	
	
	
	

	Exceeding the limits for leakage current in single fault condition.
	
	
	
	

	exceeding the voltage limits in the case of single fault condition (for basic insulation)
	
	
	
	

	Starting, interrupting or locking of movements
	
	
	
	

	
Single fault conditions


1
Overloading of mains supply transformers in equipment (test see clause 57.9)
	
	
	
	

	Failure of thermostats (see clause 52.5. 10 and 56.6)
	
	I
	
	No thermostats.

	Short circuiting of either constituent part of a double insulation
	
	
	
	

	Interruption of the protective earth conductor
	
	
	
	

	Impairment of cooling
	
	
	
	

	Locking of moving parts (see clause 52.5.8)
	
	/
	
	No moving parts.

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations; Nctes/Comments

	52.5.7
	Interruption and short circuiting of motor capacitors (see clause 52.5.8)
	
	I
	
	No motors or capacitor motors.

	
	Capacitor
and Winding
	Short
Circuit
	Open
Circuit
	
	
	
	Ri
	R2
	ti
	t2
	t
	Final Temp

	52.5.8
	Additional tests for motor operated equipment


Do the temperatures not exceed the limits of table XI and XII?
	
	I
	
	Duration of test:
mmVoltage:
vFinal temperature of winding:

	52.5.9
	Failure of components
	
	I
	
	
	
	
	
	
	

	52.5.10


~


-___
	Overload
	
	I
	
	
	
	
	
	
	

	
	a) Equipment with heating elements


1.
Thermostats (see clause 52.5. lOc and
52.5. lOd)
	
	I
	
	No heating elements.

	
	2.
Equipment with heating elements with short time rating (see clause 52.5. lOc and 52.5. lOe)
	
	I
	
	
	
	
	
	
	

	
	3. other equipment (see clause 52.5. lOc)
	
	I
	
	
	
	
	
	
	

	
	b) Equipment with motors


1.
Motor is part of equipment (see
52.5.5
through 52.5.8 and 52.5. lOf through 52.5. lOh)
	
	I
	
	
	
	
	
	
	

	
	2. Equipment containing motors as well
as heating parts
	
	I
	
	
	
	
	
	
	

	
	3.
If more than one of the tests is
applicable for the same equipment,
these tests are made consecutively
	
	I
	
	
	
	
	
	
	

	
	c) Equipment having heating elements
	
	I
	
	No heating elements.

	
	d) Heating parts of equipment
	
	I
	
	
	
	
	
	
	

	
	e) Heating parts of equipment (part 2)
	
	/
	
	
	
	
	
	
	

	
	0
Motors with overload protection
	
	I
	
	

	
	g) Equipment rated for short-time or intermittent operation
	
	I
	
	
	
	
	
	
	


	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	

	
	h) Equipment for three phase motors
	I
	1’
	
	No motors.
—_____________

	56.


56.1
	Components and General Assembly


Is the constructional data form with a list of critical parts available? (mains part and applied_part)
	
	
	
	

	56. ib)
	Marking of components


Do the ratings of components not conflict with the conditions of use?
	I
	
	
	

	56.ld)
	Component fixing


Is unwanted movement of components prevented?
	
	I
	
	Does not apply to component evaluation.

	56.10
	Fixing of wiring


Are conductors and connectors so secured that accidental detachment cannot result in a safety hazard?
	
	I
	
	Does not apply to component evaluation.

	56.3


56.3a)
	connections - General


Is incorrect interconnection of accessible connectors prevented where a hazard may be caused (electrical, hydraulic & pneumatic)?
	
	/
	
	No accessible connectors.

	
	Are accessible parts separated from live parts? (see 17g)
	
	I
	
	

	
	Can plugs for connection of patient circuit leads not be connected to other outlets?
	
	I
	
	No interconnecting plugs.

	
	Are medical gas connections not interchangeable?
	
	I
	
	No gas connections.

	56.3b)
	Are accessible conductive parts prevented from becoming live when connection between different parts of the equipment is broken?
	
	I
	
	No accessible conductive parts.

	56.4


~~__________
	Connection of capacitors
Are capacitors not connected between live parts and non-protectively earthed accessible parts?
	
	
	.
	

	
	Do capacitors connected between the mains part and protectively earthed accessible parts comply with IEC 384-14?
	
	/
	
	No capacitors across terminals.

	
	Are enclosures of capacitors provided only with basic insulation not secured to non-protectively earthed accessible parts?
	
	I
	
	No capacitors.



95

‘I

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	56.4 cont
	Are capacitors not connected between thermal cut-out terminals?
	I
	
	No capacitors.

	56.5
	Protective devices
Are protective devices to disconnect the supply mains which operates by producing a short circuit which results in operation of an overcurrent device not used?
	I
	
	No protective devices. Does not apply to component evaluation.

	56.6


56.6a)
	Temperature and overload control devices Application


Are thermal cut-outs which require soldering to reset not used?
	I
	
	No temperature or control devices. Does not apply to component evaluation.

	
	Are thermal safety devices provided to prevent operating temperatures exceeding the specified limits? (see 57.9)
	I
	
	No thermal safety devices. Does not apply to component evaluation.

	
	Is an independent non-self-resetting thermal cut-out provided where the failure of a thermostat may cause a hazard?
	I
	
	No thermal cut outs. Does not apply to component evaluation.

	
	Does the operating temperature of the above thermal cut-out lie between the upper limit of the first thermostat and the safe limit for function?
	I
	
	No thermal cut outs. Does not apply to component evaluation.

	
	Is there an audible alarm where loss of function could present a safety hazard?
	/
	
	No temperature or control devices. Does not apply to component evaluation.

	
	Are heated liquid containers protected against dangerous overheating when empty?
	I
	
	No heated liquid containers.

	56.ob)
	Is the adjustment range of thermostats not much greater than required for equipment function?
	I
	
	No temperature or control devices. Does not apply to component evaluation.

	
	Is the temperature setting clearly indicated?
	I
	
	

	
	Is the operating temperature of thermal cut-outs clearly indicated?
	I
	
	

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	56.7


56.7a)
	Intern-al electrical power source Housing


Are housings containing batteries:
Adequately ventilated?
	I
	
	No batteries.

	
	constructed to prevent accidental short circuiting of the batteries?
	I
	
	No batteries.

	56.7b)
	Connection


Is the equipment fitted with means to prevent incorrect polarity of connection?
	I
	
	No batteries.

	56.8
	Indicators


Unless indication is provided by other means, are indicator lamps used:


to indicate that equipment is energized?
	I
	
	Does not apply to component evaluation.

	
	to indicate the operation of non-luminous heaters?
	I
	
	Does not apply to component evaluation.

	
	to indicate when outputs are energized? (required only where a safety hazard could result)
	I
	
	Does not apply to component evaluation.

	
	Is a charging mode provided and is it visibly indicated to the operator?
	I
	
	Does not apply to component evaluation.


No controls. Does not apply to component evaluation.

	56.10


56. lOa)
	Actuating parts of controls


Protection against electric shock Do accessible parts of electrical controls comply with the requirements of 16c)?
	I
	
	

	56. lOb)
	Fixing, prevention of maladjustment Are all actuating parts so secured that they cannot be pulled off or work loose?
	I
	
	No controls. Does not apply to component evaluation.

	
	Are controls so secured that the indication of any scale always corresponded with the position of the control?
	I
	
	No controls. Does not apply to component evaluation.

	
	Is an incorrect connection of an indicating device prevented by an adequate construction, if it can be separated without the use of a tool?
	I
	
	No controls. Does not apply to component evaluation.



Clause
Description, Requirements

No.

Rotating Controls

Gripping diameter of

the knob [mm]

56.llb)
Mechanical strength


Do hand-held control devices comply with the


requirements of 21.5?

Are foot operated control devices able to

withstand the weight of a human being? (1.35

- Kn for 1 minute)

Are foot-operated controls drip-proof?

Are electrical switching parts of foot-operated control devices in operating rooms of watertight construction?

I

No control devices.

56. lOb)

cont.

Test Torque

[Nm]

Compliance Results, Observations, Notes/Comments 0KM/A Fail

	1O_d_23
	1.0
	
	/
	

	23_d_31
	1.8
	
	/
	

	31_d_41
	2.0
	
	I
	

	41_d_56
	4.0
	
	/
	

	56_d_70
	5.0
	
	1
	

	Axial Pull
~
	Test Force
60 N
	
	I
	


Mcd~anicai Cctnp~n~s
100 N

I

	56. lOc)
	Limitation of movement
Are steps provided to prevent an unexpected change from maximum to minimum or vice-versa and to prevent damage to wiring?
	I
	

	56.11


56.lla)
	cord connected hand held and foot-operated control devices


Limitation of operating voltage Do operating voltages of cord connected
control switches not exceed 25 v a.c. or 60 V d.c (see clause 17g)
	I
	No control devices. Does not apply to component evaluation.
Measured voltage:
V


I
No control devices.

I
No control devices.

56.1 ic) Inadvertent operation

Do control devices not change their control _________ setting in abnormal position?

56.lld) Entry of liquids

I
No control devices.

I
No control devices.

	Clause No.
	Description, Requirements
Compliance
0KM/A Fail
	Results, Observations, Notes/Comments

	56.1 le)
	Connection cords
Does the connection and anchorage of flexible cord to control devices comply with the requirements of 57.4?
	I
	
	No connection cords.

	57.
	Mains Parts, Components and Layout
	I
	
	Does not apply to component evaluation.

	57.1


57.la)
	Isolation from the supply mains Isolation


Does the equipment have means for simultaneous disconnection of all supply poles?
	
	
	

	
	Means for isolation:

incorporated in the equipment?
	I
	
	Does not apply to component evaluation.

	
	external (specified in accompanying documents)?
	I
	
	Does not apply to component evaluation.

	57. id)
	Does any switch fitted to comply with the clause 57. la) above meet the requirements of
IEC 328?
	I
	
	No switches. Does not apply to component evaluation.

	57.10
	Are mains switches not incorporated in flexible mains leads?
	I
	
	No switches. Does not apply to component evaluation.

	57. ig)
	Are directions of movement of actuators of switches in accordance with JEC 447?
	I
	
	No switches. Does not apply to component evaluation.

	57. lh
	If there is a suitable plug device to isolate a non-permanently installed equipment from the supply mains, does it comply with 571.a)?
	I
	
	No switches. Does not apply to component evaluation.

	57. im)
	Are fuses and semiconductor devices not used as isolating devices?
	I
	
	No fuses or semiconductor devices.

	57.2
	Mains connectors, appliance inlets and the like
	I
	
	Not a mains operated appliance.

	57.2e)
	Are auxiliary mains socket outlets of a type that cannot accept a mains plug? (see 56.3). Are they properly marked?
	
	
	

	57.3
	Power supply cords
	I
	
	Not a mains operated appliance.

	S7.3a)
	Is the equipment not provided with more than one connection to a particular supply mains?
	
	
	

	
	Does no hazard result if more than one connection is made simultaneously?
	I
	
	Not a mains operated appliance.

	
	Is the mains plug not fitted with more than one power supply cord?
	I
	
	Not a mains operated appliance.



 Clause
Description, Requirements
Compliance
Results, Observations, Notes/Comments


No.

0KM/A Fail


57.3q)
Is the equipment provided with a power
I
No supply cord. Does not apply to


cont
supply cord or with an appliance inlet?

component evaluation.


57.3b)
T)pes
I
No supply cord. Does not apply to





component evaluation.



Do power supply cords comply with IEC 245



designation 53 or with IEC 227 designation



53?


Are polyvinyl chloride insulated power supply
I
No supply cord. Does not apply to


cords not used where external metal parts with

component evaluation.


a temperature exceeding 75°C exist?

    -


3c)
Cross-sectional area of conductors
I
Area:
min2Cu

Do supply conductors comply with the cross

_________ sectional areas in table XV?

57.3d)
Preparation of conductors
I
No supply cord. Does not apply to


Are stranded conductors of cables fixed by

component evaluation.

_________
any clamping means but not soldered?


57.4
Connection of Power Supply Cords
I
No supply cord. Does not apply to



Chord anchorages

component evaluation.

Does equipment and mains connectors provided with power supply cords have cord anchorages?


Are the cord anchorages made of:
/
No supply cord. Does not apply to




component evaluation.


1.
Insulating material  OR -


2.
Metal insulated by supplementary
/



insulation from accessible conductive



parts non-protectively earthed?


3.
Metal provided with an insulated
I



lining?

Do clamping screws not bear directly on the
I


cord insulation?

Do screws operated when replacing the power
I supply cord not serve to fix any component other than parts of the cord anchorage?

Are conductors of the power supply cord so
I arranged, that if the cord anchorage fails, the protective earth conductor is not subject to


strain?
-
I I

	Clause
Description, Requirements
	Compliance 0KM/A Fail
	

	Cord guards
For other than stationary equipment, is the flexible supply cable adequately protected against excessive bending?
	
	I
	
	No supply cord. Does not apply to component evaluation.

	Is an opening in equipment so shaped that the applied supply cord (even if not provided with guards) passes the flexing test (clause 57.4b)?
	
	
	
	No supply cord. Does not apply to component evaluation.

	Accessibility of the connection Is there sufficient space inside the equipment
to allow the supply cable conductors to be introduced and connected?
	
	I
	
	Does not apply to component evaluation.

	Can any covers be fitted without risk of damage to the conductors or their insulation?
	
	I
	
	Does not apply to component evaluation.

	Is it possible to check that conductors are correctly connected and positioned before the cover is fitted?
	
	I
	
	Does not apply to component evaluation.

	Mains Terminal Devices and Wiring of Mains Part


Is mains connected equipment other than that fitted with a detachable supply cord provided with mains terminals where connections are made with screws, nuts etc.?
	
	I
	
	Not a mains connected appliance. Does not apply to component evaluation.

	Is a barrier provided to prevent a reduction of the values of creepage distances and air clearances between live parts and other conductive parts, should the conductor break away? (see 57.10)
	
	I
	
	Not a mains connected appliance. Does not apply to component evaluation.

	Are the screws and nuts not used to secure other components (other than internal conductors likely to be displaced)?
	
	I
	
	Does not apply to component evaluation.

	Arrangement of maths terminal devices Are terminals including any protective earth terminal closely grouped as to allow easy connection?
Protective earth conductor, see clause 58. Marking of mains terminal, see clause 6.2.
	
	I
	
	Not a mains connected appliance. Does not apply to component evaluation.

	Are they inaccessible without the use of a
tool?
	
	
	
	Not a mains connected appliance. Does not apply to component evaluation.


Mains Fuses and Overcurrent Releases

List of fuses and overcurrent releases:

Device, type and rating

Are there fuses or overcurrent releases in each

supply lead for class I equipment and for class

Ii
equipment having a functional earth

(according to 18.1)?

I

Does not apply to component evaluation.

	Clause No.
	Description, Requirements
	Compliance
0KM/A Fail
	

	57.5b)
continued
	Are mains terminal devices so located or shielded that, should a wire of a stranded conductor escape when the conductors are fitted, there is no risk of accidental contact between live parts and accessible parts and, for class II equipment, between live parts and conductive parts separated from accessible parts by supplementary insulation only?
	I
	
	

	57.5c)
	Fixing of mains terminals


Does tightening or loosening of clamping means of conductor not:


1.
Subject internal wiring to stress?
	I
	
	Not a maths connected appliance. Does not apply to component evaluation.

	
	2.
Reduce creepage distances and air
clearances below allowable limits?
	I
	
	Not a mains connected appliance. Does not apply to

	
	Cross sectional area of the conductor used in the test:
	I
	
	component evaluation.


~— ~°Cu

	57.5d)
	Connections to mains terminals


Does effective connection to mains terminals not require special preparation of conductors (also see clause 57.3d)?
	I
	
	Not a mains connected appliance. Does not apply to component evaluation.

	
	Are conductors not damaged or displaced by tightening or loosening of clamping screws or nuts?
	I
	
	Does not apply to component evaluation.


‘~7. 6

I
Does not apply to component evaluation.


Location in the circuitry

Maths supply transformers

List of mains supply transformers:

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Conunents

	57.6 cont
	Is the current rating of maths fuses and overcurrent releases such that they reliably carry the normal operating current and not greater then the current rating of any component in the mains circuit carrying the mains supply current?
	
	I
	
	Does not apply to component evaluation.


Does

	
	Is the protective earth conductor not fused?
	
	
	
	

	
	Is no fuse fitted in the neutral conductor of permanently installed equipment?
	
	I
	
	not apply to component evaluation. Does not apply to component evaluation.


No mains wiring. Does not apply to component evaluation.

	57.8


.... I. 8a)


57.8b)
	Wiring of mains part


Insulation


If insulation of an individual conductor is not at least equivalent to that required by IEC 227 or IEC 245, that conductor is considered as bare.
	
	I
	
	

	
	Cross section


Is the cross sectional area of wiring between the maths terminal and fuse sufficient to prevent any fire hazard in case of fault currents (see clause 57.3c)?
	
	
	
	Measured area: ——
mni2Cu

	
	Is the cross section of other wiring and PC board tracks in the maths part adequate to prevent a fire hazard in the event of a fault?
	
	
	
	


57.9


57.9
Type of transformer:

continued

57.9.1 Overheating

Are mains supply transformers protected against overheating of basic insulation, supplementary insulation and reinforced insulation in the event of short circuit or overload on any output winding?

Are protective devices external to the transformer or its enclosure provided against overheating connected in such a way that failure of any component other than wiring interposed between the protective device and the transformer cannot render the protective device inoperative?

Test under conditions specified in clause 42.

until steady thermal conditions are obtained.

e)
Oyercurrent releases

1)
Protective device not provided for limitation of the winding

temperature?

Compliance Results, Observations, Notes/Comments 0KM/A Fail


/
What current was used as the basis for the


overload test?
A


I
If the secondary outputs are shorted in turn,



does a fuse rupture? If so which fuse



ruptured? or

Is there another means for prevention of



overheating?

If none of the above, the trans6Drmer must

be considered a failure

I

57.9.lb)
Overload

	S7.9.la)
	Short circuit
	I

	
	Conditions
	specified in clause 42.
	



I
Does not apply to component evaluation.

	Loading of a section or each winding in turn of the transformer:
	I
	Does not apply to component evaluation.

	a)
With fuses which comply with IEC 127 and IEC 241?
	I
	

	b)
With fuses which deviate from IEC 127 and JEC 241?
	/
	

	c)
Short circuit current is smaller than test current?
	I
	

	d)
Thermal cut-outs?
	I
	


I

I

NOTES ON FAILURES

Report No. 556395

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	57.9b)
cont
	Determination of the temperature rise of Transformer copper windings by the resistance method:


At=(R2-R1/R1)x(2345+tl)+(t241)


Overall compliance with clause 42?
	I
	
	Insulation class:
Room temperature: ti =
Room temperature: t2 =
t=
OC

	R1W1=
	1=
	A
	R2W1=
	=~T=
	

	R1W2=
	1=
	A
	R2W2=
	°~T=
	

	R1W3=
	1=
	A
	R2W3=
	=~T=
	

	R1W4r=
	1=
	A
	R2W4=
	~T=
	°C

	R1W5=
	1=
	A
	R2W5=
	=T=
	°C

	R1W6=
	1=
	A
	R2W6=
	=~T=
	°C

	R1W7=
	1=
	A
	R2W7=
	=~T=
	°C

	RIW8=
	1=
	
	
	
	


‘ ~‘

S7.9.4e)
Does reinforced or double insulation between


the primary and secondary winding consist of:

One insulation layer having a thickness of at least 1 mm? OR

Two insulation layers with a total thickness of

not less than 0.3 mm? OR

Three layers provided that each combination of two layers can withstand the dielectric strength test for reinforced insulation?

I

No transformers. No transformer

No transformers.

I

	Clause No.
	Description, Requirements
	Compliance OKN/AFajl
	Results, Observations, Notes/Comments


Does not apply to component evaluation. No transformers.

	57.9.2
	Dielectric Strength
Are the requirements of clause 20 fulfilled?
	
	I
/
	
	

	
	(if yes, they shall not be repeated!)
	
	
	
	

	
	Do electrical insulations between turns and layers pass the test of 57.9.2 after the humidity pre-conditioning treatment?
	
	I
	
	Test voltage:
VTest

	57.9.4 S7.9.4a)
	Construction Methods used for separation:


Separate bobbins or formers?
	
	I
	
	
frequency:
Hz


No transformers.

	
	One bobbin or former with imperforate partition between windings?
	
	I
	
	No transformers.

	
	One bobbin or former with concentric windings and having an imperforate protective copper screen with a thickness not less than 0.13 mm?
	
	I
	
	No transformers.

	
	One bobbin with concentric windthgs separated by double or reinforced insulation?
	
	I
	
	No transformers.

	57.9.4c)
	Are means provided to prevent displacement of end turns?
	
	I
	
	No transformers.

	57.9.4d)
	Protective earthed screen with one turn:
	
	I
	
	No transformers.

	
	Does the screen overlap at least 3 mm?
	
	I
	
	No

	
	Is the width of the screen at least equal to the axial winding length of the primary winding?
	
	I
	
	
transformers.


No transformers.


I

No transformers.

Do creepage distances between the primary and secondary windings comply with the requirements for reinforced insulation?

I

	Clause No.
	Description, Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes!Comments

	57.9.4g)
	Are the exit of the wires from the internal windings of toroidal transformers provided with double sleeving complying with the requirements of double insulation?
	
	I
	
	No transformers.

	
	Is the total thickness at least 0.3 mm, extending at least 20 mm outside the winding?
	
	I
	
	Measured thickness:
mmMeasured length:
mmNo transformers.

	57.10
57.lOa)
~
	Creepage distances and air clearances


Do creepage distances and air clearances comply with the values in table XVI?
	
	I
	
	Component produces a 4-20 millivolt potential. Creepage and clearance distances are adequate for intended application.

	
	Are creepage distances for slot insulation of motors at least 50% of the values of table XVIwithamjnimumof2n~at~OV?
	
	I
	
	Component produces a 4-20 millivolt potential.

.

	
	For insulation in the mains part between parts of opposite polarity the minimum creepage distances and air clearances are not required if short-circuit of the creepage distances and air clearances in turn does not produce a safety hazard.
	
	I
	
	Not a mains supplied component.

	Insulation between
	Reference voltage
[VJ
	
	I
	
	Air clearance [mm]
	Creepage distance [mm]

	Basic ~ bctwcc~p5tt,o( cppc~itc po1in~y


Buicot
‘~~‘~Y
issuistion


~
	A-f
	
	—
	
	
	Required
	Measured
	Required
	Measured

	
	A-al
	
	
	
	—
	
	
	
	

	
	A-b
	
	
	
	
	
	
	
	

	
	A-c
	
	
	—
	
	
	
	
	

	
	A-j
	
	
	
	
	
	
	
	

	
	B-c
	
	
	
	
	
	
	
	

	
	B-d
	
	
	
	
	
	
	
	

	°~
Rcurforcc4
UtMrlaIjcs,
	A-a2
	
	
	
	
	
	
	
	

	
	A-e
	
	
	
	
	
	
	
	

	
	A-k
	
	
	
	
	
	
	
	

	
	B-a
	
	
	-
-
I
	-
	
	
	
	

	~-
	B-e
Overall compliance with S7. lOa)?
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	


	Clause No.
	Description,
	Requirements
	Compliance 0KM/A Fail
	Results, Observations, Notes/Comments

	58.
	Protective earthing
	-Terminals and
	
	I
	
	

	
	connections
	
	
	
	
	No protective earthing required, Not a

	58.1
	
	
	
	
	
	maths operated component.

	
	Does the protective means comply with 57.5c)?
	earth terminal clamping the requirements of
	
	
	
	No protective earthing required.

	
	Is it not possible to of a tool?
	loosen it without the use
	
	/
	
	

	
	Are screws for internal connections completely against inadvertent
	protective earth covered or protected
loosening from the outside
	
	I
	
	No protective earthing required.

	58.7
	of the equipment?
	
	
	
	
	

	
	Is the earth pin of the as the protective inlet forms the supply
	appliance inlet regarded terminal, where an appliance
connection to
	
	I
	
	No protective earthing required. Not a mains operated component.

	58.8
	equipment?
	
	
	
	
	

	
	Is the protective earth the mechanical parts of the equipment component not related
	terminal not used for connection between different
or the fixing of any to protective earthing
	
	I
	
	No protective earthing required. Not a mains operated component.

	58.9
	or functional earthing?
	
	
	
	

	
	Protective earth connection

~
	
	I
	
	No protective earthing required. Not a

	
	Is the protective earth and interrupted after are made or interrupted, between mains supply equipment is made via
	connection made before the supply connections
where the connection conductors and
a plug and socket
	
	
	
	mains operated component.

	59.
	device? (see also 57.1,
	57.2
& 57.3)
	
	I
	
	

	
	Construction and layout
	
	
	
	Does not apply to component evaluation.

	59.1
	Internal wiring
	
	
	
	
	

	
	Mechanical protection
	
	
	
	

	
	Are internal cables and
	wires protected from:
	
	
	
	

	
	1.
Contact with2.
Friction with
	moving parts? sharp edges, burs or
	
	/
	
	

	
	corners?
	
	
	
	
	Does not apply to component evaluation.

	
	3.
Is wiring having provided with there is movement parts where it them?
	
basic insulation
additional sleeving, if relative to metal
is in direct contact with
	
	I
	
	Does not apply to component evaluation.

	
	
	
	
	
	


Are internal cables and wires and components not likely to be damaged by opening or assembling equipment?

Bending

Is the bending radius of cables and cable forms at least 5 times the outer diameter of the lead?

Can insulating sleeving only be removed by breaking or cutting or is it secured at both ends?

Are aluminum wires of less than 16 mm2 cross section not used?

Applicable requirements

Are connecting cords between equipment parts (e.g. parts of an X-ray patient monitoring installation or a data processing installation or combinations) considered belonging to equipment and not be subject to requirements for wiring of electrical installations?

Insulation

Mechanical strength and resistance to heat and fire

Does insulation comply with the requirement

Moisture resistance? (see clause 44) Dielectric strength? (see clause 21)

The following tests need not be carried out where evidence of compliance is provided:

1.
Ball pressure test?

2.
Do parts of insulation material supporting

parts at mains voltage satisfy the requirements

of the ball pressure test at a temperature of

125°C

Does not apply to component evaluation.

Ball pressure test temp: — Does not apply to component evaluation.

Does not apply to component evaluation.

58.1 cont

Description, Requirements

Date:
11/08/95

59.lb)

I

Compliance
Results, Observations, Notes/Comments

0KM/A Fail

Does not apply to component evaluation.

59.lc)

I

Insulation

Does not apply to component evaluation.

/

No sleeving.

	Does
	the sheath of a flexible cord comply
	I
	

	with
	the requirements of IEC 227 and IEC
	
	No flexible cords.

	245,
	if it is used as supplementary insulation?
	
	


59. id)

Do insulated conductors, which are subject to temperatures greater than 7S°C, have an insulation of heat resistant material?

Materials

I

Does not apply to component evaluation.

I

No aluminum wires.

I

Does not apply to component evaluation.

59.10

59.2

59.2b)

of:

Does not apply to component evaluation.
Does not apply to component evaluation.

	Clause No.
	Description, Requirements
Compliance
0KM/A Fail
	Results, Observations, Notes/Comments

	59.2c)
~
	Protection


Is insulation not likely to be impaired by deposition of dirt or dust resulting from wear of parts within the equipment?
	I
	
	Does not apply to component evaluation.

	
	Are ceramic materials, not tightly sintered, not used as supplementary or reinforced insulation?
	I
	
	No ceramic materials.

	
	Are rubber materials used as supplementary insulation in class II equipment resistant to ageing? (oxygen test 59.2c)
	I
	
	No rubber materials used for supplementary insulation.

	
	Are creepage distances not reduced below those specified in 56.10 despite any cracks in such insulation?
	I
	
	Does not apply to component evaluation.

	
	Are insulation materials in which heating conductors are embedded not used as reinforced insulation?
	I
	
	No heating conductors.

	59.3
	Excessive current and voltage protection Has an internal electrical power source in equipment an appropriately rated device to protect against fire hazards?
(test under consideration)
	I
	
	Does not apply to component evaluation.

	
	Are fuse elements replaceable without opening the enclosure fully enclosed in a fuseholder?
	I
— I
	
	No fuses. Does not apply to component evaluation.

	
	Are live parts of fuseholders shielded to prevent electric shock when replacing fuses? (if replaceable without the use of a tool)
	
	
	No fuseholders. Does not apply to component evaluation.

	
	Do protective devices between isolated applied part and the enclosure not operate below
500 V RMS?
	I
	
	No protective devices. Does not apply to component evaluation.

	59.4
	Oil containers
Are oil containers in portable equipment adequately sealed to prevent oil loss in any position and does the oil container design allow for the expansion of oil?
	/
	
	No oil containers.

	
	Are oil containers in mobile equipment sealed to prevent loss of oil during transport? (a pressure release device may be fitted)
	I
	
	No oil containers.

	
	Is an oil level indicator provided on partially sealed oil filled components?
	I
	
	No oil containers.


IEC601.doc                                 Page 10 of 33

