VANDAGRAPH
  COMPANY OPERATING PROCEDURE

Oxygen Sensor manufacture

Modification of R-22DEM Sensors 

VM3/COP/40.14

Date: 28 Jun 2007
Revision Date: 28-Aug-09
Issue 1

Only to be attempted by trained personel




Parts Required

Quantity
Description
Part No.


1
R-22D Sensor (any version will work R-22D is preferred R-22A is lowest cost and has a grid)



1
White Box 
9040021



White Box  Lid
9040002


2
Packing Material
9040006


1
Round “Vandagraph” Label
9040003


2
Serial number Labels



1
Serial number Label for Outer bag



1
Polythene Bag (outer)



1
PCB Blank or




PCB pre-wired with resistors
9015005


2
4K7 ohm surface mounted resistors or 



2
5K ohm surface mounted resistors



1
Cable Red/White wires + low profile 2 way Molex
9070120


[image: image1.jpg]


[image: image2.jpg]



1. Open the packaging and remove the R-22D 


Sensor from the Gas barrier bag

2. Check the O2 sensor for damage and signs of 

leaking electrolyte.

3. Remove the O2 sensor specification leaflet and 

discard.  

4. [image: image3.jpg]


[image: image4.jpg]


Remove excess outer case unto level of PCB. 

VM3COP40.11

5. Using a Flow divertor hold the sensor in a vice

6. Remove the Molex carefully. The PCB is non   

conformal coated so the seal must be broken on 

the sensor carefully 

7. W[image: image5.jpg]


[image: image6.jpg]


here possible use Pt number  9015005                   
pre-wired PCB with 4K7 or 5K6 resistors and molex 
connected  If not available see page 2

8. Test for output 4.0 mV- 6.5mV.

9. Wire so that the leads lie naturally across the PCB

10. Cut the three pins as close to the solder joint as 

possible.

11. Remove the existing label and add R-22DEM label to sensor  & Tube so that the edges of the labels are parallel to the edges of the container and Sensor.

12. Ensure the O2 sensor is booked out in the stock books. Place packing material in the bottom and top of the white container and add the cap. Ensure it is sealed.

13. Add the Vandagraph round label to the top of the cap. Add the instruction leaflet, and place the whole into the polythene bag, 





Notes for PCB construction

14. [image: image7.jpg]


[image: image8.jpg])




[image: image9.jpg]


The PCB should be cleaned with a PCB rubber to remove oxidation.

15. The Red & White wires should be cut to 5 cms and stripped with the tool. Insert the wire until the end touches the Red stop. Close the tool and the cable will be stripped. If not rotate the wire and repeat.

16. Tin the bare ends.

17. Inserts should be crimped onto the wires using the special tool 

18. The 4K7 or 5K6 resistors should be soldered onto the PCB using the Microscope set up. Only to be attempted by trained SMT engineer. Use two resistors of the same value on the same PCB

4581%5B1%5D 090828.doc       28/08/2009      Page 2 of 2

