Comparison of Viamed DB3 Supramaximal Nerve Stimulator V MIPM TOF3D Neuromuscular Transmission Monitor

|  |  | Viamed Microstim DB3 | MIPM TOF3D |  |
| --- | --- | --- | --- | --- |
|  |  | (subjective) | (subjective) | (objective) 3-dimensional acceleromyography |
| Stimulation modes |  |  |  |  |
| TOF - Train of Four | Single |  |  |  |
|  | Continuous, every 15 seconds. |  |  |  |
| TOFs - Train of Four | Stimulation interval adjustable between 1 to 60 minutes. |  | Stimulation mode can be removed/added via the setup menu. |  |
| PTC - Post Tetanic Count |  |  |  |  |
| Single Twitch stimulation | Single | 1 Hz | Default can be switched between 1 Hz and 0.1 Hz Stimulation mode can be removed/added via the setup menu. |  |
|  | Continuous | 1 Hz | Default can be switched between 1 Hz and 0.1 Hz Stimulation mode can be removed/added via the setup menu. |  |
| DBS - Double Burst Stimulation |  | Ratio 3:2 | Default can be switched between ratio of 3:2 and 3:3 Stimulation mode can be removed/added via the setup menu. |  |
| TET - Tetanic |  |  | Default can be switched between 50 Hz and 100 Hz Stimulation mode can be removed/added via the setup menu. |  |

Page 1 of 4

| Stimulation sites |  |  |
| --- | --- | --- |
| Abductor pollicis / nervus ulnaris |  |  |
| Flexor hallucis brevis muscle / posterior tibial nerve |  |  |
| Obiculari oculi muscle / Facial nerve |  |  |

| Stimulation features & options |  |  |
| --- | --- | --- |
| Stimulation strength unit | mA | Default can be switched between mA or µC |
| Stimulation pulse width | 200 µs | In any stimulation mode the pulse width can be switched between 200 µs or 300 µs |
| Stimulation output | Rotary control knob | Adjustable via keypad. Default output can also be set (0 - 60 mA, 0 - 12/18 µC) |
| Numerical display of stimulation output |  |  |
| Stimulation visual indicator |  |  |
| Stimulation beep |  | Can be switched on/off |
| Timer display |  |  |

| Calibration modes |  |  |  |
| --- | --- | --- | --- |
| Automatic calibration function for alignment with patients’ individual response |  |  |  |
| Impedance monitoring for optimal stimulation |  |  |  |
| Calibration of device and creation of baseline reference |  |  |  |
| Calibration (1) at set stimulation strength |  |  |  |
| Calibration (2) determination of supramaximal stimulation strength |  |  |  |
| Adjustable baseline |  |  |  |

Page 2 of 4

| Graphs & trends |  |  |  |
| --- | --- | --- | --- |
| Single Twitch Trending Graph |  |  |  |
| TOF Trending Graph |  |  |  |
| TOF Twitch Result graph |  |  |  |
| PTC result graph |  |  |  |
| Single Twitch stimulation graph |  |  |  |
| TOF stimulation graph |  |  |  |
| DBS stimulation graph |  |  |  |
| TET stimulation graph |  |  |  |

| Display features |  |  |  |
| --- | --- | --- | --- |
| Large LCD display |  | 4.4" (11.18 cm) |  |
| Adjustable TOF monitor alerts (lower & upper) |  |  |  |
| Alerts and faults display |  |  |  |
| Skin Temperature monitoring |  | Sensor attaches to the hand adapter. |  |

| Patient stimulation data |  |  |
| --- | --- | --- |
| Onboard data storage |  |  |
| Data storage download to external devices |  | (Unreleased) |
| Remote connection to patient monitor |  | (Unreleased) |

| Accessories |  |  |  |
| --- | --- | --- | --- |
| Stimulation leads | Included | Included |  |
| Patient cable assembly |  | Included |  |
| AMG (Acceleromyography) sensor |  |  | Included |
| Temperature sensor |  | Optional accessory |  |
| AMG Thumb adapter |  |  | Optional accessory |
| AMG Hand adapter |  | Optional accessory, for connection of skin temperature sensor. | Optional accessory, for connection of AMG sensor and/or skin temperature sensor. |
| AMG Eye adapter |  |  | Optional accessory |
| Pole mount / standard rail mount |  | Optional accessory |  |
| USB cable |  | Optional accessory |  |

Page 3 of 4

| Stimulation Electrodes | Uses standard low-cost ECG electrodes | Uses standard low-cost ECG electrodes (Paediatric) |
| --- | --- | --- |

| Designed to comply with the latest European medical standards |  |  |
| --- | --- | --- |

Hz 	= 	Stimulation pulses per seconds (Hertz).

mA 	= 	Stimulation strength in milliamperes.

µC 	= 	Stimulation strength in microcoulombs.

µs 	= 	Stimulation pulse width in microseconds.

VIAMED

Page 4 of 4

V1.1      2021-07-15