Tek-Ox Serial Number TX0003

Non Auto switch off version ( Green switch)

Fitted with new batteries 07/08/08 and left in the ON position

10/10/08 Lo Batt shows and the unit closes down within 24 hrs. showing a blank screen
If batteries removed and re-inserted
Lo Batt shows
Calibration in 100% Oxygen
Allowed to settle in Air 20.9% Reading 20.8%

John S. Lamb

10/10/08