	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Conversion R-17D-VG to R-17D-VGJJ

	VM3/COP/40.48

	Date: 28-07-09
	Revision Date: 6-Oct-09
	Issue 1

	Only to be attempted by trained personel





	Parts Required

	Quantity
	Description
	Part No.
	

	1
	R17D-VG
	
	

	1
	4mm Grommet
	9040100
	

	1
	1cm Heatshrink
	9080018
	

	1
	0.5cm Heatshrink
	
	

	1
	Oertling cable clamps
	9040201
	

	1
	Connector RS 295-5609 to be supplied with cable
	
	

	1
	Cable   RS 521-8206 to be supplied with connector
	
	

	1
	White Box 
	9040021
	

	
	White Box  Lid
	9040002
	

	2
	Packing Material
	9040006
	

	1
	Round “Vandagraph” Label
	9040003
	

	2
	Serial number Labels
	
	

	1
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	


[image: image1.jpg]


[image: image2.jpg]


[image: image3.jpg]


Procedure

1. Remove the top cap from the R-17D-VG, Cut the through the label carefully with a scalpel.

2. [image: image4.jpg]


[image: image5.jpg]


Prise of the cap and remove the jack socket




3. Remove the red and Black wires. Do not apply too much heat to the PCB.

4. [image: image6.jpg]


Insert the grommet and glue into position








5. Feed the cable through the grommet, strip the cable , separate the screen ,add heat shrink over the screen and a small piece of heat shrink at the junction of the signal wire and the screen. 
The Oertling clamp should be about 2 cms from the end of the cable

6. Solder the wires to the hole/pads Red to Red Black to Black.




7. [image: image7.jpg]


Re-assemble and test. Output should be 9mV to 13mV.Glue the cap into position 






8. Remove existing label and replace with an R17D-VGJJ
[image: image8.jpg]












9. Package[image: image9.jpg]


[image: image10.jpg]


[image: image11.jpg]



6180%5B1%5D.doc       06/10/2009      Page 2 of 2

