QC 33b


	[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





	
	[image: image2.png]


 Ltd.
15 Station Road, Cross Hills, Keighley, West Yorkshire, BD20 7DT, UK.

Website: www.viamed.co.uk.  Email: info@viamed.co.uk.
Tel: +44 (0)1535 634542.  Fax: +44 (0)1535 635582.


TOM THUMB CALIBRATION / TEST & Q.A. SHEET.

	Description.
	Tom Thumb.

	Model.
	TT 490-15. - Restricted

	Serial No.
	

	Time & Date of Calibration / Test.
	

	Time & Date of QA.
	


Do not start the QA check within 1 hour of the calibration / test.

Labelling is to be attached after calibration / test.

All Tests are performed at 10 Lpm with a restrictor in place.

Record the manometer reading in millibars below, ensuring the TT490-15 meets the limits specified.

	Test Equipment
	Test
	Specification
	Reading
	P/F

	
	
	
	Cal
	QA
	

	
	
	
	
	
	

	Snoop liquid.
	Check all ports / connections for leaks.
	No bubbling.
	√
	√
	P

	
	
	
	
	
	

	CE 149
	Adjustable Valve:

@ 10 Lpm
	Minimum:

( 8 cmH2O.
	
	
	P

	
	
	
	
	
	

	CE 149
	Adjustable Valve:

@ 10 Lpm.
	Tolerance:

( 43 & ( 47 cmH2O.
	
	
	P

	
	
	
	
	
	

	CE 149
	Precision Valve:

@ 10 Lpm.
	Maximum:

+3.0 cmH2O over the adjustable valve setting.
	
	
	P

	
	
	
	
	
	

	CE 149
	Pressure Gauge Test.

@ 50 cmH2O.
	Tolerance: ± 2% of

Precision Valve Reading
	√
	√
	P

	
	
	
	
	
	

	Visual check.
	Flowmeter reading remains steady after shake test
	
	√
	√
	P

	
	
	
	
	
	

	Visual check.
	Gauge cannot be removed without tools.
	
	√
	√
	P

	
	
	
	
	
	

	Visual check.


	Gauge appears straight.
	
	√
	√
	P

	
	
	
	
	
	

	Visual check.
	All adjustable settings are set to a minimum.
	
	√
	√
	P

	
	
	
	
	
	

	Visual check.
	Labels are attached.
	CE label.
	√
	√
	P

	
	
	Serial No. label.
	√
	√
	

	
	
	O2 Flowmeter Label
	√
	√
	

	
	
	Tom Thumb label
	√
	√
	


Calibration:
E. Sendrowski



Q.A Check: 

Signed:  …………………………………                   

Signed:  ………………………………






0

