	VIAMED COMPANY OPERATING PROCEDURE

	Calibration of the DL-3000 SpO2 Simulator.

	VM3/COP/45.03

	Date: 2002
	Revision Date:22-Sep-06
	Issue: 2


DL-3000 SpO2 Simulator Operating Instructions.

    Software Version : 10528M1E. 

Issue 2.

Software Version : 10528M1E.

The DL-3000 only requires a simple calibration procedure which should be carried out annually, to guarantee the accuracy of the unit.  

The DL-3000 calibration procedure involves circuitry adjustments as detailed below.  Test equipment used (DVM) should be calibrated and traceable to National Standards.  Viamed Ltd. cannot be liable for the accuracy of the unit once the original calibration seal is broken.

1. Remove the rear panel retaining bolts (4) and separate the rear panel from the remainder of the casing.  Slide out the upper casing cover.

2. Remove the shorting link at ‘Lk2’ from the RUN position and refit in the CAL position.

3. Identify the Issue No. printed on the motherboard under the test finger connector. 

“Iss.A” : Carry out steps (1 - 10).  “Iss.B”, “Iss.C” : Carry out steps (1 - 7) and (10).

4. Switch on the DL-3000 whilst F4 is depressed.  The unit enters its calibration routine.  Whilst F4 is depressed the unit displays it’s software version. 

	
	
	Trim vr2 for zero
	

	
	
	tp11 - to - tp12                     next
	

	
	
	
	
	
	
	
	
	
	

	(
	
	F1
	
	F2
	
	F3
	
	F4
	

	(
	
	
	
	
	
	
	
	
	

	Enter : No function.


5.  Connect a calibrated DVM between Tp11 & Tp12.  Trim Vr2 until the voltage between Tp11 and Tp12 equals 0.00V.  Press F4.
Selections.
F1, F2 & F3 : No function.

F4 : Select next measurement.

	
	
	Trim vr3 for zero
	

	
	
	tp11 - to - tp12                     next
	

	
	
	
	
	
	
	
	
	
	

	(
	
	F1
	
	F2
	
	F3
	
	F4
	

	(
	
	
	
	
	
	
	
	
	

	Enter : No function.


6.  Trim Vr3 until the voltage between Tp11 & Tp12 equals 0.00V.  Press F4.
Selections.
F1, F2 & F3 : No function.

F4 : Select next measurement.

	
	
	Trim vr1 for zero
	

	
	
	tp11 - to - tp12                     next
	

	
	
	
	
	
	
	
	
	
	

	(
	
	F1
	
	F2
	
	F3
	
	F4
	

	(
	
	
	
	
	
	
	
	
	

	Enter : No function.


7.  Trim Vr1 until the voltage between Tp11 & Tp12 equals 0.00V.
Selections.
F1, F2 & F3 : No function.

F4 : Select next measurement.

This completes the necessary adjustments for Iss.B and Iss.C motherboards.  Go to (10).
Steps (8) & (9) for “Iss.A” motherboards only.
	
	
	with probe connected
	

	
	
	meas tp13 – tp14                done
	

	
	
	
	
	
	
	
	
	
	

	(
	
	F1
	
	F2
	
	F3
	
	F4
	

	(
	
	
	
	
	
	
	
	
	

	Enter : No function.


8.  Connect the DVM between Tp13 and Tp14.  Measure and record the voltage difference.
Selections.
F1, F2 & F3 : No function.

F4 : Select next measurement.

	
	
	with probe connected
	

	
	
	trim vr4 for prev
	

	
	
	
	
	
	
	
	
	
	

	(
	
	F1
	
	F2
	
	F3
	
	F4
	

	(
	
	
	
	
	
	
	
	
	

	Enter : No function.


9.  Trim Vr4 until the voltage between Tp13 and Tp14 equals the recorded value at (8).
Selections.
F1, F2 & F3 : No function.

F4 : Return to (5).

10.  Switch off the unit, refit ‘Lk2’ in the RUN position and reassemble the unit.

The unit must be switched off to exit the calibration routine.
DL-30 & DL-300 : Calibrated Light Source & Test Finger Calibrator.  

The DL-30 is a calibrated light source used to verify the DL-300 test finger calibrator.  The instrument produces a precise infrared light output of 45µW.  The DL-300 is a calibrated test unit which connects to the test finger and verifies the correct output of light using an independent photo detector.  The calibration is traceable to National Standards.







VM3COP 45.03 DL3000 Calibration.doc                              22/09/2006                         Page 0 of 2
1
1

