	CERTIFICATE  OF CALIBRATION

Issued by:                   Vandagraph  Ltd.

Date issued:               9 March 2010
Certificate No.:            

Signed:                       ____________




Client 

Date of calibration. 
9 March 2010
Instrument 
Mixcheck


Serial number: 



Procedure I.D.  

VM3COP40.56



Performance Status  
PASS

The equipment detailed herein has been calibrated  and now conforms to the manufacturer’s published specifications. The Q.A. of the above is in strict conformity to BS EN ISO 9001-2000 & ISO 13485-2003.  All measurements made are traceable to NAMAS.

	Equipment used to carry out calibration

	Fluke Graphical Multimeter, serial number 6911002, CE080
	

	Pico Technology Virtual Instrument ADC-200, serial number 3468, CE020
	

	Microcal  Serial Number
	

	Oxygen BOC Medical 99.5% License 07351
	

	Helium  BOC Medical 99.5% License 07351
	


	Setting
	Oxygen Reading
	Helium Reading

	100% Oxygen
	
	*

	Dry Air
	
	

	100% Helium
	
	

	Trimix
	
	


* 100% Oxygen will produce a small error of approximately 1.4% (helium) due to the difference in the Thermal conductivity of Air and Oxygen.

Vandagraph Cal Cert Mixcheck.doc                            9th March 2010


