Use of VM-2101 / VM-2105 with SPO2 Simulator Index 2

Simulator: Index 2
Manufacturer: Fluke / Bio-Tek

Software: 3.00
Curve/Make: BCI

AMP: 3.000%

TLC 1100
Ambient Normal

	Index 2
BCI Curve
Pulse 75
Saturation [ % ]
	Reading on VM-2101/VM-2105
[ % ]

	
	min
	typical
	max

	100
	99
	99
	99

	99
	99
	99
	99

	98
	99
	99
	99

	97
	98
	98
	98

	96
	97
	97
	97

	95
	96
	96
	96

	94
	95
	95
	95

	93
	94
	94
	94

	92
	92
	93
	93

	91
	91
	92
	92

	90
	91
	91
	91

	89
	90
	90
	90

	88
	89
	89
	89

	87
	88
	88
	88

	86
	86
	87
	87

	85
	85
	86
	87

	84
	84
	85
	85

	83
	83
	84
	84

	82
	83
	83
	84

	81
	81
	82
	82

	80
	80
	81
	82

	79
	79
	80
	81

	78
	78
	79
	79

	77
	77
	78
	78

	76
	76
	77
	77

	75
	75
	76
	77

	74
	74
	75
	75

	73
	73
	74
	74

	72
	72
	73
	73

	71
	71
	72
	72

	70
	70
	71
	72


[image: image1.emf]

Please note that according to the international standard ISO9919:2005,

chapter 50.101.2 the use of patient simulators or functional testers shall 

not be used to validate the accuracy of pulse oximeter equipment!
The only reliable method of testing the measurement accuracy of an SpO2 monitor is the clinical validation of the measurement data, indicated by the monitor on the basis of blood gas analysis.

During extensive clinical studies, the VM-2101 and the VM-2105 evidenced the accuracy required.

09/03/10
