	     [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley - West Yorkshire  BD20 7DT - United Kingdom

Tel: +44 (0)1535 634542         Fax: +44 (0)1535 635582

Email: info@viamed.co.uk       www.viamed-online.com



Validation of Flowsensor A with Dräger Ventilators and Monitors
[image: image1.png]Hot-wire flowsensor to measure gas flow in breathing systems.

Compatible to the Dräger Spirolog sensor 84 03 735.

Flowsensor A Part Numbers
4310001 (pack of 5)

4330001 (single unit quantity)

Typical Compatibility Test Results

One Dräger Spirolog reference flowsensor was used to compare against five standard Flowsensor A sensors, test date 11 June 2008.
MV = Minute volume, l/min

Dräger Cato Ventilator with Divan
	Setting

Vt/l
	Setting

f 1/min
	Dräger

Spirolog Vt/MV
	Viamed

No. 1

Vt/MV
	Viamed

No. 2

Vt/MV
	Viamed

No. 3

Vt/MV
	Viamed

No. 4

Vt/MV
	Viamed

No. 5

Vt/MV

	1
	10
	0.99/9.9
	0.97/9.8
	0.98/9.9
	0.99/1.0
	0.98/9.9
	0.98/9.8

	1
	20
	0.99/19.7
	0.97/19.5
	0.97/19.4
	0.99/19.8
	0.98/19.6
	0.99/19.7


Dräger Evita 4 ICU Ventilator
	Setting

Vt/l
	Setting

f 1/min
	Dräger

Spirolog Vt/MV
	Viamed

No. 1

Vt/MV
	Viamed

No. 2

Vt/MV
	Viamed

No. 3

Vt/MV
	Viamed

No. 4

Vt/MV
	Viamed

No. 5

Vt/MV

	1
	10
	0.99/9.9
	0.99/9.8
	0.99/9.9
	0.99/9.9
	0.99/9.8
	0.99/9.9


Dräger EVA ICU Ventilator
	Setting

Vt/l
	Setting

f 1/min
	Dräger

Spirolog Vt/MV
	Viamed

No. 1

Vt/MV
	Viamed

No. 2

Vt/MV
	Viamed

No. 3

Vt/MV
	Viamed

No. 4

Vt/MV
	Viamed

No. 5

Vt/MV

	1
	10
	0.97/9.6
	0.97/9.8
	0.98/9.9
	0.99/9.8
	0.97/9.6
	0.98/9.7


Dräger Anemone Monitor
	Setting

Vt/l
	Setting

f 1/min
	Dräger

Spirolog Vt/MV
	Viamed

No. 1

Vt/MV
	Viamed

No. 2

Vt/MV
	Viamed

No. 3

Vt/MV
	Viamed

No. 4

Vt/MV
	Viamed

No. 5

Vt/MV

	1
	10
	0.97/9.6
	0.98/9.7
	0.98/9.8
	0.99/9.8
	0.96/9.6
	0.97/9.7


Date: 03/10
[image: image2.jpg]


