	[image: image3.jpg]



	VIAMED
	V1.7

Part No. 2590014                            


	Technical Manual

Microstim DB3

Supramaximal Nerve Stimulator
Peripheral nerve stimulator 

[image: image4.jpg]


for use during anaesthesia


CE0086

Table of contents

Section 1 

Cautions & warnings – page 2.
Section 2 

Cleaning instructions – page 2.
Section 3  
Packaging & return information – page 3.
Section 4

Warranty information – page 4.
Section 5 

Equipment required to test the Microstim DB3 – page 4.
Section 6 

Routine testing procedure – page 4.
Section 7 

Disassembly of the Microstim DB3 – page 8.
Section 8 

Removal of PCB – page 9.
Section 9 

Fault finding guide – page 10.
Section 10 
Test procedure – Output & battery level indicator (LED) – page 11.
Section 11
Test procedure – Audio transducer – page 12.
Section 12
Test procedure – Output signal – page 13.
Section 13
Test procedure – Intensity control potentiometer – page 14.
Section 14
Test procedure – Mode switch 1 – page 15.
Section 15
Test procedure – Mode switch 2 – page 16.
Section 16
 Parts list – page 17.
Section 17   Manufacturer’s details – page 17.
Section 1 – Cautions & warnings

[image: image5.jpg]



      Caution:

A prerequisite before using this manual is that the user should be familiar with the operation of the Microstim DB3 in accordance with the operating instructions supplied with the unit.
[image: image6.jpg]


      Caution:
The circuitry inside the Viamed Microstim DB3 does not require any periodic maintenance or calibration.  There are no user-serviceable parts.  Viamed does not recommend field repair of the Viamed Microstim DB3. Any sign or evidence of opening the unit or field service by non-Viamed personnel; tampering, or any kind of misuse or abuse of the system shall void the warranty in its entirety.  All repairs and service should be performed by Viamed, or a service representative authorized by Viamed.
The Microstim DB3 is designed to be robust yet compact in size. Without access to specialised equipment, troubleshooting and repair of circuit board components is not feasible.

[image: image7.jpg]



      Caution:

Dispose of used batteries in accordance with local ordinances and regulations.

[image: image8.jpg]



     Do not use a patient cable that appears to be damaged.
[image: image9.jpg]



There are no user-serviceable parts inside the Microstim DB3. The casing should only be removed by qualified service personnel.

[image: image10.jpg]


Pulse voltages in excess of 90V DC may be encountered on the PCB assembly.
Section 2 – Cleaning instructions
To surface clean the Microstim DB3 and patient cable:

Use a soft cloth dampened with either a commercial non-abrasive cleaner, or a solution of 70% alcohol in water.

Lightly wipe the surfaces of the unit.

To disinfect the Microstim DB3 and patient cable:

Use a soft cloth saturated with a solution of 10% chlorine bleach in tap water. 

Section 3 - Packaging and return procedure

Obtain a Return Authorization number and shipping instructions before returning a DB3 for service by contacting the manufacturer / service centre:

Viamed Ltd.

15 Station Road

Cross Hills

Keighley

West Yorkshire BD20 7DT

United Kingdom

Tel. +44 (0)1535 634542

Fax. +44 (0)1535 635582

Email: info@viamed.co.uk
Please clean contaminated equipment before returning it to Viamed. 

Follow the cleaning procedures as in Section 2 - do not autoclave. 

Package the equipment in the original packaging materials. If the original packaging is not available, use a suitable carton with approprioate packing material to protect it during shipping.

Enclose a letter specifying the following items:

1. Return Authorization number obtained from Viamed.

2. Service required:

For repair service please quote reference - 2540000. 
Detail any difficulties experienced with the Microstim DB3.

 
For routine annual service please quote reference - 2580000.

3. 
The serial number of the Microstim DBS. The serial number label is located inside the battery compartment.
4. Shipping and invoicing information of the sender for returning the serviced Microstim DB3 and invoicing for any chargeable repair costs.

5. Contact information for any questions about the service; including:


name, department and email/telephone/fax number.
Return equipment by a method that provides proof of delivery.

Section 4- Warranty information 

This product is warranted against defects in material and workmanship, and to operate within published specifications under normal use for a period of one year from the date of original shipment. Batteries are not warranted.

To request a warranty claim, contact Viamed for return authorization and instructions.

If an examination by Viamed discloses such products or component parts to be defective, Viamed’s sole obligation is limited to repair or replacement (at Viamed’s option) of the defective product or component and shall not include shipping costs.

This warranty does not extend to any product that was subject to misuse, neglect or accident; that was damaged by causes external to the product; or that was used in violation of the operating instructions supplied with the product. This warranty does not extend to any product that was modified in any way, or disassembled or reassembled by anyone other than Viamed or an authorized Viamed agent. 

This warranty, together with any other express written warranty that may be issued by Viamed is the sole and exclusive warranty as to Viamed’s products. This warranty is expressly in lieu of any oral or implied warranties; including without limitation any implied warranty of merchantability or fitness for a particular purpose. Viamed shall not be liable for any incidental, special or consequential loss, damage or expense directly or indirectly arising from the loss of use of any products.

Section 5 - Equipment required to test the Microstim DB3

1) Digital multi-meter.

2) DB3 Tester, part number 2510005.

[image: image11.emf]
3) Oscilloscope capable of displaying 4MHz pulses, with a vertical graticule 0.1 Volts/division, and Horizontal sweep of 200 micro seconds/division minimum.

Section 6 - Routine testing of the Microstim DB3
a) Testing the battery. Using a digital multi-meter, test the voltage of the battery.

[image: image12.jpg]


Ensure that the battery is a 9V alkaline battery such as the following:

   Duracell Procell MN1604

b) Testing of the patient lead. Test the continuity of the individual leads of the patient lead assembly using a digital multi-meter.

The standard patient cable assembly supplied with the Microstim DB3 is:
[image: image13.emf]
Part number 2520000
Length: 0.90 m

Terminations: Press-studs:

Alternative patient cables available are as follows:

Part number 2520001
Length: 1.5 m

[image: image14.jpg]


Terminations: Press-studs

Part number 2520003
Length: 1.5 m

Terminations: Miniature pinch clamps

c) Routine testing procedure.
i)   Ensure that the battery of the Microstim DB3 is in good condition. Without the patient lead connected press one of the function buttons, (in this case P.T.C is better). If the battery condition is good then the Microstim DB3 LED indicator will be green. The LED indicator changes from green to amber, then red as the battery output decreases. Replace the battery if the LED indicator is red.

ii) Connect the Microstim DB3 Tester to the output of the Microstim DB3, red (+ve) and black (-ve).  Set the Microstim DB3 intensity control to maximum (fully clockwise).

[image: image1.jpg]



iii) Depress each of the Microstim DB3 function buttons in turn.  When a function button is pressed and held there should be audio output in conjunction with illumination of the green LED indicator on the Microstim DB3 (refer to the following Modes of Stimulation). The output can be verified by observing if the red LED indicator on the Microstim DB3 Tester is also illuminated in synchronization. 

Modes of Stimulation:

Train of Four (T.O.F)

Consists of four stimuli at a frequency of 2 Hz.

[image: image15.jpg]POST TETANIC COUNT

| 5s | 3s—11sq
Tetanic Recovery Single Twitches
50 Hz 1 Hz, Continuous




Double Burst Stimulation (D.B.S)

Consists of two short bursts of stimuli (3:2) at 50 pulses per second separated by 
750 ms between bursts.

[image: image16.jpg]DOUBLE BURST STIMULATION

|-20 ms-| |20 ms
| 750 ms |





Post Tetanic Count (P.T.C)

Consists of a sequence of pulses; 50 Hz for 5 seconds, a 3 second pause, followed by 1 Hz stimuli.

[image: image17.jpg]TRAIN OF FOUR





Continuous 1Hz (Single Twitch)

Consists of a single stimulus applied repetitively at a frequency of 1 Hz.

[image: image18.jpg]



Diagrams not to scale; for illustration purposes only. 

[image: image19.jpg]


Note: Stimuli pulse width 200μs. To observe the stimuli pulses the DB3 tester needs to be connected to an oscilloscope – refer to section e).

d) Test results:

i) Microstim DB3 Tester LED indicator illuminated red – Microstim DB3 

output test passed. It indicates that a stimulus of at least 90 V is being produced and that the battery voltage is at least 7.5 V.
ii) Microstim DB3 Tester LED indicator not illuminated:

· Battery level maybe low, check battery condition 


as in c)i) above.

· If still not illuminated check intensity control set to maximum 


as in c)ii) above.

· If still not illuminated then the Microstim DB3 has developed a fault,


i.e. no or low output.

e) Verification of the output using an oscilloscope.

In order to further verify the output of the unit, additional tests can be carried out by connecting the Microstim DB3 Tester to an oscilloscope using the supplied BNC cable supplied with the tester.

Connect one end of the BNC cable to the output of the Microstim DB3 Tester and the other end to a suitable oscilloscope.

[image: image2.jpg]



[image: image20.jpg]@ [Microstim[[=K}

Supramaximal Nerve Stimulator

A (2]

ATTENTION, SEE INSTRUCTIONS FOR USE

Caution when used with cardisc pacemakers
Roplace battery when indicator flashes red

C€0086

REF| 2510000

Al __2009.04

== = 3





The Microstim DB3 Tester reduces the output of the Microstim DB3 by a factor of 100 in order to facilitate easy testing.

Typically the oscilloscope can be set to amplitude of 0.1 V per division and the time base as appropriate to the output being observed, i.e. 200 μs per division.
Section 7 – Disassembly of the Microstim DB3
[image: image21.jpg]



[image: image22.jpg]



[image: image23.jpg]



[image: image24.jpg]



Remove the battery compartment cover and the four retaining screws.

Open the case carefully. The case is held at the top with two retaining hooks; care should be taken as these can be broken. Prise case apart carefully with fingers until the retaining hooks are disengaged. 


Once the two retaining hooks have been disengaged the case can be opened and the two halves laid side by side. 


The intensity control knob can now be pulled off the spindle of the control potentiometer. 


   Reassembly notes:

a) Ensure that the intensity control knob is correctly aligned, i.e. turn the potentiometer fully counter clockwise and seat the knob in the spindle locating notch - so that the arrow on the intensity control knob indicates minimum intensity.
b) Fit the two halves of the case together by first securing the retaining hooks. When fitting the two halves of the case together ensure that the battery leads are not trapped.
Section 8 – Removal of PCB assembly. 
Stage 1:

Desolder the two mode switches; both switches have six solder pads.
Desolder the intensity potentiometer – three solder pads.

Remove the two PCB securing screws.

                             Intensity potentiometer


Switch 1



Switch 2


                              PCB securing screws.

The PCB can now be pulled away from the case; but is still connected to the battery compartment, input connector and the audio transducer. Stage 1 allows for the replacement of the mode switches or the intensity potentiometer.

Stage 2:
Desolder the battery connections, output connector and audio transducer. This enables the whole PCB assembly to be removed, including the LED indicator.

                                                                       Output connector
Battery 

connections                                                    Audio transducer
Section 9 - Fault finding guide
	DB3
	DB3 Tester
	Oscilloscope
	Action

	Mode switches
	Audio
	LED
	Output control
	Red LED illuminated
	Test range of output
	

	All four modes


	Yes
	Green
	Set to maximum
	Yes
	Pass
	Non required - unit functioning correctly

	All four modes


	Yes
	Red
	Set to maximum
	
	
	Ensure intensity control set to maximum.

If DB3 LED still red – replace battery.

	All four modes


	Yes
	No illumination
	Set to maximum
	Yes
	Pass
	Test LED – page 11.

	All four modes


	No 
	Green
	Set to maximum
	Yes
	Pass
	Ensure DB3 tester connected. 
DB3 requires a ‘patient load’ for audio to function. 
If still no audio then test speaker – page 12.

	All four modes


	Yes
	Yes
	Set to maximum
	No
	None
	Test output – page 13.

	All four modes


	Yes
	Yes
	Test span
	Yes
	Fail
	Test potentiometer – page 14.

	T.O.F.


	No
	No
	Set to maximum
	No
	None
	Test mode switch 1 – page 15.

	D.B.S.


	No
	No
	Set to maximum
	No
	None
	Test mode switch 1 – page 15.

	P.T.C.


	No
	No
	Set to maximum
	No
	None
	Test mode switch 2 – page 16.

	1 Hz


	No
	No
	Set to maximum
	No
	None
	Test mode switch 2 – page 16.


10

Section 10 - Test procedure - Output & battery level indicator (LED):
Disassemble unit as per instructions on page 8.
Using an oscilloscope determine if the LED drive signal is present. 

Select 1 Hz mode during test.

Using a multi meter carry out diode test i.e. ≈1.6 V drop across the green diode and ≈1.7 V across the red diode.

                                                                        Red anode

                                                                        Cathode

                                                                        Green anode

	LED drive signal present?
	LED  test
	Action
	Part number

	Yes
	Fail
	*Replace LED, (pcb ref: LED1).


	2530016

	No
	Pass
	**Replace pcb sub assembly.


	2500000


*Necessitates removing the PCB assembly as per procedure stage 1, see page 9.

**Necessitates removing the PCB assembly as per procedure stage 1 and 2, see page 9.


Part number

2500000
11

Section 11 - Test procedure – Audio transducer:

Disassemble unit as per instructions on page 8.
Using an oscilloscope determine if the audio transducer drive signal (+5 V) is present. 
Select 1 Hz  mode during test.


                                                                        -ve
                                                                       Audio transducer

                                                                      +5 V
                                                                                                            (Capacitance 22 nf +/- 30%)
	Audio transducer drive

signal present?
	Action
	Part number

	Yes


	*Replace audio transducer
 (pcb ref: B1).
	2530014

	No
	**Replace pcb sub assembly.


	2500000


*Necessitates removing the PCB assembly as per procedure stage 1, see page 9.

**Necessitates removing the PCB assembly as per procedure stage 1 and 2, see page 9.


Part number

2500000
Section 12 - Test procedure - Output signal

Disassemble unit as per instructions on page 8.
Using an oscilloscope determine if output signal is present. Set oscilloscope to 20V/div.

Select  DB3 1 Hz  mode during test.

                                                                       Output connector
                                                                       Output -ve


                                                                       Output +ve

	Output present?


	Action
	Result
	Part number

	Yes
	Test continuity of output between pcb and connector output pins.


	Open circuit on +ve or –ve? *If so, replace output connector assembly.
	2530010

	No
	Test intensity control potentiometer, see next procedure.


	If potentiometer OK **replace pcb sub assembly
	2500000


*Necessitates removing the PCB assembly as per procedure stage 1, see page 9.

**Necessitates removing the PCB assembly as per procedure stage 1 and 2, see page 9.


Part number

2500000
Section 13 - Test procedure  - Intensity control potentiometer

Disassemble unit as per instructions on page 8.
Using an oscilloscope determine if stimulation signal is present. Set oscilloscope to 2V/div.

Select  DB3 1 Hz  mode during test.

Centre pin – stimulation signal from the micro controller.

Right hand side pin – output to power amplification stage.

                             Intensity potentiometer


                              .                                       

	Output present?


	Test potentiometer (10K) span using digital multi meter
	Action
	Part number

	Yes
	Fail
	*Replace potentiometer
	2530013

	No
	Pass
	***Replace pcb assembly
	2500000


*Necessitates removing the PCB assembly as per procedure stage 1, see page 9.

**Necessitates removing the PCB assembly as per procedure stage 1 and 2, see page 9.


Part number

2500000
Section 14 - Test procedure - Mode Switch 1

Disassemble unit as per instructions on page 8.
Using a digital multi-meter carry out voltage measurements.



TP 3                                                              TP6
TP 1                                                              TP 4


TP 2                                                              TP 5

Battery +ve
Battery  -ve
	Mode switch selection
	Test point
	Expected voltage
	Result
	Action
	Part number

	Non 
	TP 1
	1) 9V (battery). 
	0V
	**Replace PCB sub assembly
	2500000

	
	TP 4
	5V regulated. 
	0V
	**Replace PCB sub assembly
	2500000

	T.O.F. 
	TP 2 & 3
	1) 9V (battery).
	0V
	*Replace switch assembly SW1
	2530011

	
	TP 5
	5V regulated.
	0V
	*Replace switch assembly SW1
	2530011

	D.B.S.
	TP 2 & 3
	1) 9V (battery).
	0V
	*Replace switch assembly SW1
	2530011

	
	TP 6
	5V regulated.
	0V
	*Replace switch assembly SW1
	2530011


1) Battery voltage minus voltage drop across polarity protection diode circuit.

*Necessitates removing the PCB assembly as per procedure stage 1 - see page 9.

**Necessitates removing the PCB assembly as per procedure stage 1 and 2 - see page 9.

NOTE: Faulty switch can be confirmed by carrying out continuity test using digital multi-meter.


Part number

2500000
Section 15 - Test procedure - Mode Switch 2

Disassemble unit as per instructions on page 8.
Using a digital multi- meter carry out voltage measurements.


TP 3                                                              TP6

TP 1                                                              TP 4


TP 2                                                              TP 5

	Mode switch selection
	Test point
	Expected voltage
	Result
	Action
	Part number

	Non 
	TP 1
	1) 9V (battery). 
	0V
	**Replace PCB sub assembly
	2500000

	
	TP 4
	5V regulated. 
	0V
	**Replace PCB sub assembly
	2500000

	P.T.C.
	TP 2 & 3
	1) 9V (battery).
	0V
	*Replace switch assembly SW2
	2530011

	
	TP 5
	5V regulated.
	0V
	*Replace switch assembly SW2
	2530011

	1 Hz
	TP 2 & 3
	1) 9V (battery).
	0V
	*Replace switch assembly SW2
	2530011

	
	TP 6
	 5V regulated.
	0V
	*Replace switch assembly SW2
	2530011


1) Battery voltage minus voltage drop across polarity protection diode circuit.

*Necessitates removing the PCB assembly as per procedure stage 1 – see page 9.

**Necessitates removing the PCB assembly as per procedure stage 1 and 2 – see page 9.

NOTE: Faulty switch can be confirmed by carrying out continuity test using digital multi-meter.


Part number
2500000
Section 16 – Parts list
	Part number


	Description

	2500000
	PCB sub assembly

	2520000
	Patient lead – standard. 0.9m, press-studs. 

	2520001
	Patient lead. 1.5m, press-studs.

	2520003
	Patient lead. 1.5m, miniature pinch clamps

	2530001
	Upper case

	2530002
	Lower case

	2530003
	Lower case label

	2530004
	Battery cover

	2530005
	Battery contact, +ve

	2530006
	Battery contact, -ve

	2530007
	Cable, battery contacts

	2530008
	Case screw

	2530009
	PCB assembly screw

	2530010
	Input connector

	2530011
	Mode switch & bezel (pcb ref. SW1 & SW2)

	2530012
	Intensity control knob 

	2530013
	Potentiometer (pcb ref. W1)

	2530014
	Audio transducer

	2530016
	LED

	2590000
	Operator’s manual

	9950055
	Battery, 9V Alkaline, MN1604


Section 17 - Manufacturer

Viamed Ltd.

15 Station Road
Cross Hills

Keighley

West Yorkshire BD20 7DT

United Kingdom

Tel:  +44 (0)1535 634542


Fax: +44 (0)1535 635582

Email: info@viamed.co.uk


Website: www.viamed.co.uk

Quality standards accreditations:

BS EN ISO 9001-2000

ISO 13485-2003

CMDCAS

Note:  Specifications subject to change. 

            Please contact Viamed for further specifications and information. 05/04/2010
         








PAGE  
1


