Instructions for use: 

Series 400-AC Temperature Probes

Contains: 1 piece
Applies to models ref:

2009/2010/2011/2012/2013/2014/2015/2016/2017/2018/2019/2020/2021/2022/2023/2024/2025/2026/2027/2028/2029/2030/2031/2032   

See overleaf for details.
Proper use

The Series 400-AC reusable temperature probes are intended for the continuous temperature measurement and monitoring of patients when used with YSI, Hewlett Packard/Philips, Siemens and other manufacturers’ patient monitors and devices, that are specially intended for use with YSI 400 series temperature probes and compatibles. They are available in various configurations of sizes and types.

Instructions for use

Temperatures can be taken by placing the temperature probe on the skin surface, in the oesophagus, or rectally using the appropriate model.

Skin

1. Position the probe with a metal side to the skin.

2. Secure the probe using a heat reflective pad, taking care not to compress blood vessels.
3. Secure the lead using medical adhesive tape. 
4. The temperature probe requires a three to five minute stabilising time.
Oesophageal
1. Introduce the probe until it has reached the lower third of the oesophagus. 
2. Secure the lead using medical adhesive tape.
3. The temperature probe requires a three to five minute stabilising time.
Rectal

1. Place a disposable rectal probe cover over the end of the temperature.
2. Lubricate the end of the probe.
3. Insert the probe into the rectum.
4. The probe is in the correct position when it has passed the sphincter.
5. Secure the lead using medical adhesive tape.
6. The temperature probe requires a three to five minute stabilising time.
Cleaning and disinfecting
1. Clean and disinfect or autoclave before using the temperature probe on another patient.

2. Before cleaning or disinfecting the probe, disconnect it from the monitoring device.

3. Clean the probe of any adhesive tape and bioburden material prior to disinfection or sterilisation.

4. Temperature probes can be cleaned with a cleaning agent by wiping the probe with a soft cloth.

5. To wipe clean, hold the lead near the probe end and wipe towards the plug.
6. To disinfect, wipe the probe and contact surfaces with disinfectant. For this we recommend the use of CIdex® OPA. 
7. If you intend to use a glutaric aldehyde or an alternative product, please follow the usage instructions for these products. 

8. Rinse thoroughly with water.
Steam sterilisation (Autoclaving)
1. Clean and/or disinfect the temperature probe prior to sterilisation.
2. Steam autoclave to 134°C using standard cycles, taking note of any warnings or precautions provided with the autoclave.
Specifications 

Reusable 

Autoclavable to +134 °C 
Storage temperature: -20°C to +70°C

Operating temperature: +0°C to +50°C

Accuracy: In the range of +25°C to +50°C ± 0.1°C

Warning

1. This temperature probe is intended for use only in conjunction with compatible monitoring devices. Usage in conjunction with other devices can lead to malfunction.

2. Strong sources of electromagnetic fields, e.g. electro-surgical devices, can adversely affect the functioning of the equipment. If at all possible, the probe should be completely removed from the patient before a surgical device or other HF source is activated! Check the protective earthing if you intend to use HF devices and temperature measuring equipment in parallel. Minimise the potential danger by selecting a temperature measurement site distal to the HF current path.

3. We do not recommend the use of the product within the range of Magnetic Resonance Imaging (MRI) devices.

4. Stop using the probe immediately if it becomes damaged.

5. For further information and warnings please read the usage instructions accompanying the temperature measuring device.

Caution

Improper use of temperature probes can lead to damage of the internal wiring and to loss of electrical insulation, as well as incorrect temperature readings.

Returning the product

Contact the customer service department at Viamed Ltd.  to obtain a RMA number (Return Material Authorization). Every product sent to us for checking must be cleaned and disinfected beforehand.
[image: image11.jpg]VIAMED



[image: image12.wmf]0086


[image: image13.wmf]0086

[image: image14.jpg]



	Part number
	Model ref.
	Description
	Connector

	0212009
	2009
	Oesophageal/rectal – neonatal, ( 2mm
	6.35mm 

(¼") jack (mono)
	[image: image1.jpg]




	0212010
	2010
	Oesophageal/rectal – paediatric, ( 3mm
	
	

	0212011
	2011
	Oesophageal/rectal – adult, ( 4mm
	
	

	0212012
	2012
	Oesophageal/rectal – neonatal, ( 2mm
	6.35mm 

(¼") jack (mono)

Right angled
	[image: image2.jpg]




	0212013
	2013
	Oesophageal/rectal – paediatric, ( 3mm
	
	

	0212014
	2014
	Oesophageal/rectal – adult, ( 4mm
	
	

	0212015
	2015
	Oesophageal/rectal – neonatal, ( 2mm
	Siemens
	[image: image3.jpg]




	0212016
	2016
	Oesophageal/rectal – paediatric, ( 3mm
	
	

	0212017
	2017
	Oesophageal/rectal – adult, ( 4mm
	
	

	0212018
	2018
	Oesophageal/rectal – neonatal, ( 2mm
	Hewlett Packard
	[image: image4.jpg]




	0212019
	2019
	Oesophageal/rectal – paediatric, ( 3mm
	
	

	0212020
	2020
	Oesophageal/rectal – adult, ( 4mm
	
	

	0212021
	2021
	Skin contact – neonatal, ( 6mm
	6.35mm 

(¼") jack (mono)
	[image: image5.jpg]




	0212022
	2022
	Skin contact – paediatric, ( 10mm
	
	

	0212023
	2023
	Skin contact – adult, ( 16mm
	
	

	0212024
	2024
	Skin contact – neonatal, ( 6mm
	6.35mm 

(¼") jack (mono)

Right angled
	[image: image6.jpg]




	0212025
	2025
	Skin contact – paediatric, ( 10mm
	
	

	0212026
	2026
	Skin contact – adult, ( 16mm
	
	

	0212027
	2027
	Skin contact – neonatal, ( 6mm
	Siemens
	[image: image7.jpg]




	0212028
	2028
	Skin contact – paediatric, ( 10mm
	
	

	0212029
	2029
	Skin contact – adult, ( 16mm
	
	

	0212030
	2030
	Skin contact – neonatal, ( 6mm
	Hewlett Packard
	[image: image8.jpg]




	0212031
	2031
	Skin contact – paediatric, ( 10mm
	
	

	0212032
	2032
	Skin contact – adult, ( 16mm
	
	



Oesophageal/Rectal


Skin 
[image: image15.jpg]


[image: image16.jpg]


[image: image9.jpg]




[image: image10.jpg]



� EMBED CorelDraw.Graphic.8  ���





Neonatal





Paediatric





Adult





Adult





Paediatric





Neonatal








[image: image11.jpg]
VIAMED Ltd.






15 Station Road, Cross Hills, 
Keighley, West Yorkshire BD20 7DT, United Kingdom 

Tel:    +44 (0)1535 634542 Fax:   +44 (0)1535 635582

Email: info@viamed.co.uk www.viamed-online.com
400-AC Temperature Probes User Instructions SW 20071008


_1250407480.unknown

