[image: image1.jpg]


VANDAGRAPH

O2 Analyser

VN202 Construction / Assembly

VM3/COP/40.05

	Date: 23 Sep. 03
	Revision Date: 17-Jun-10
	Issue: 2



“A”   Preparation of the Case


1.0 Unwrap the cases and separate them into their component parts, being careful not to lose the case screw, which is inside the battery compartment, and the window, which is inside the case.

[image: image2.jpg]


[image: image3.jpg]



2.0 With a strong sharp blade, cut the two mounting pegs nearest to the top on the back case, down by approximately 1 – 2mm.NB If it is cut down too much the PCB will short onto the EMC coating

“B”   Screening the Cases


1.0 Wrap masking tape around the sides of the back case, so that about 2mm of tape can be folded into the inside edge of the groove that runs around the case.

2.0 Cut away the piece of tape covering the bottom of the case, which forms the battery compartment.

3.0 Wrap tape around the front case, covering the outside, but not the inside, of the ridge, which runs around it. Wrap tape around the edge of the battery door.

4.0 Place all the masked parts down on a suitable surface with their inside upwards for spraying.

5.0 Place the top panels down with their largest face upwards as this will become the inside surface.

6.0 With the screening spray, cover the exposed surfaces from every angle. Allow approximately 15 minutes for everything to dry, then repeat the spraying process for a further coat. When fully dry, remove the tape. Be careful not to over spray as this may interfere with the sealing of the case. 

“C”   Drilling the Cases[image: image4.jpg]





1.0 Stick the top labels to the un-sprayed surfaces of the top panel.

2.0 Using the former and a 6mm drill bit, drill a hole between the Vandagraph Logo and the address, below the centre of the panel. Use a former



3.0 NB The front label can be added at this stage instead of section D as t[image: image5.jpg]


his gives a clean hole and no swarf to get under the label 

4.0 Using the  former and a 2.5mm drill bit drill locating holes for Calibration & ON/OFF components.

5.0 With a M131 10mm Roto Bore , drill a hole in the front case exactly in the middle, from left to right and 14mm from the top of the label area.Use a forme

6.0 With a 12mm Roto Bore  drill a hole in the front case below the first one, exactly in the middle, form left to right, and 14mm from the bottom of the label area. Use a former.

7.0 NB The Drill setup has a locating systen on the base for easy line-up.

“D”   Final Preparation of the Case

[image: image6.jpg]



1.0 Remove any dust, and drilling swarf, from the front case and carefully stick on a “Front Label”, ensuring that there are no bubbles or wrinkles on the label.

2.0 From the inside of the case cut around the holes 

3.0 With a scalpel blade. On the larger hole, cut a triangular notch about 2mm deep. 


4.0 A small groove has to be filed in the ON/Off switch hole to allow the switch through the front panel, 

5.0 Ensure the O ring is in postion on the switch ( part of the sealing kit) Place the  small bar onto the back of the switch and screw it in.

6.0 Put the rubber washer onto the potentiometer and put into the front case.  Add the crimped washer and the nut onto the potentiometer and tighten, ensuring that when viewed from the front, the white dot on the knob moves equal distances clockwise and anti-clockwise from the middle position. [image: image7.jpg]



7.0 Push a grommet through the top panel from the front making a neat fit, if necessary, by widening the hole slightly with a reaming tool. 

8.0 Squirt a small amount of silicone rubber into the channel moulded into the Rear case top, and smooth it flat. 

9.0 Fit the top panel into the silconed grove.

10.0 Join the top panel to the rear with copper tape


11.0 [image: image8.jpg]


[image: image9.jpg]


Put a battery door into the back case, and then stick a small piece of Black PVC tape or masking tape over the area where the battery door catch can be seen. Cover the tape with silicone rubber .

12.0 Ensure the battery door fits correctly and can be removed easily.




13.0 Fasten a short length of wire , bared & splayed to the rear of the case [image: image10.jpg]


with metalised tape. So that the other end can be soldered on to the -ve battery post. 

“E”   P.C.B. Assembly


1.0 The P.C.B. requires assembling using the principle that components of the lowest height are place in first. Polarity of IC’s, the Diode D1 and the Capacitor C2 must be observed.

2.0 Insert solder posts into Battery holes, Sensor holes , Switch holes, & Potentiometer holes

3.0 I  Insert link wire L, ensuring it is flat to the P.C.B.

4.0 Insert all resistors.

5.0 Insert Diode observing correct polarity.

6.0 [image: image11.jpg]


Insert IC 2 & IC 3 ensuring that they are correctly aligned, by observing the markings on the P.C.B. 

7.0 NB solder pins 234567 together (top left to right).

8.0 Insert the 28 way socket ensuring that they are correctly aligned, by observing the markings on the P.C.B.

9.0 Insert C1 & C3 – polarity is not important.

10.0 Insert C2 noting that it must be inserted with the negatively marked leg to the Outside of the P.C.B.so as to connect to the negative track on the other side of the board. As indicated on PCB.

11.0 Insert FET1 so that the flat side is to the outside of the P.C.B. As indicated on PCB

12.0 Insert P1 with the adjustment screw to the right hand side. As indicated on PCB

13.0 Cut wires on the battery connector so there is no spare and insert the battery connector ensuring Red is to + and black is to -.

14.0 Take a ribbon cable approximately 7cm long, strip and tin both ends.

15.0 Solder 2 pieces into the P.C.B. at position marked SW. Solder the other 3 pieces into the P.C.B. at the position marked VR1.

16.0 Place production stickers inside the front case, one between the switches and one inside the battery compartment, these are sequential numbers use to ensure that there is no tampering of the serial number at a later date. These are listed and linked the the serial number and LCD batch number. 
NB From S/N 4402 the PCB has the same S/N as the Instrument 

“F”   P.C.B. Spraying


1.0 The P.C.B. requires spraying with a conformal coating. 
Always use in a well- ventilated area. 
Always follow the Instructions on the can. 

2.0 [image: image12.jpg]


Spray the back of the board with conformal coating. Allow to dry. Turn the P.C.B. over and spray the other side.

3.0 Add an LCD and test the PCB on the test rig. Needs Pads 2 shorted

G”   Attaching the Coiled Cable.


1.0 Fold the dual plugged sensor cable in half and cut into two equal lengths.

2.0 Using the special tool Strip about 3cm of the outer sleeving, being careful not to damage the shield or the inner conductors. Separate the shield from the wires, Cut the shield off as close to the insulation as possible.

3.0 Strip about 3 – 4 mm of insulation from the inner conductors. Carefully tin the inner conductors.

4.0 Push the cable through the grommet of a pre-prepared end panel.  Add small O ring BS006 onto cable

5.0 Screw the PCB to the back, leaving the screw nearest the battery connector until last

6.0 Connect the wire going to the metal tape to the battery -ve

7.0 Solder the black wire to - ve post & white wire onto the + ve post.

8.0 Superglue the cable into the grommet

9.0 Superglue the O ring onto the cable near to the grommet

10.0 With a soldering iron, on the back of the panel meter, join the two halves of the solder pad marked “2”. Insert into 28-way connector ensuring that the side of the DPM with only 10 pins is nearest to the top of the board. 

11.0 solder the 3 wires from the board at the point marked”VR1” to the potentiometer on the front case. They must be connected so that the wire nearest the top of the board goes to the far tag on the pot, the middle wire to the middle tag and the bottom wire to the nearest tag.

12.0 Solder the two wires from the board at the point marked “SW” to the switch on the front case.

13.0 Add the battery connector

“H”   Testing the instrument

1.0 Connect the battery connector to a variable power supply set at 9v. Press the switch on the front case to turn the unit on.

2.0 Connect the Jack Plug to a millivolt source, switch off so as to give Zero output. Set the pot on the front case to fully clockwise. Adjust the Trim-pot “P1” until a reading of 00.0 is observed.

3.0 Set the millivolt source to 50mV. Adjust the pot on the front case to calibrate the monitor to 100.0. The reading should not drift.

4.0 Adjust the millivolt source until the LCD reads 88.8, to check the integrity of the LCD segments.

5.0 Adjust the millivolt source down to 10.5mV, the LCD should read 21.0  0.2. Switch the millivolt source off, checking that the display reads 00.0, adjusting the Trim-pot if necessary.

6.0 Add Loctite

7.0 Reduce the voltage from the power supply until the low battery indicator can just be seen to appear, the display should not drift by more than 1.0%. Turn the unit off and back on again. The time elapsed before the auto shut-off is activated should be approximately 3 minutes.

8.0 This test can be carried out with a battery

[image: image13.jpg]


“J”   Encapsulating the P.C.B.


1.0 Cut a strip of conductive foil tape approximately 1.5 cm square and fix a splayed bared wire to the front panel. The other end must reach the -ve battery post. Ensure that it is firmly stuck down the inside of the casing .It can be fixed with Silicone. NB this wire be an extra core in the ribbon cable that connects the switch to the PCB 

2.0 Put the battery door on the back of the case and connect a battery. Place the battery in the battery door so that the connector is on the right hand side of the case.

3.0 Pull the battery wires tight and tuck them into the channel moulded into the case. Put a washer onto the last screw, separate the RED and BLACK wires so that they to either side of the screw hole and insert the last screw, be careful not to over-tighten. 

4.0 The length of battery wires should be sufficient to reach the battery posts without tension but not left too long

“K”   Sealing the Instrument

1.0 Retest the instrument as before (section “H” 3.0 – 8.0), adjusting the trim-pot back to 00.0 as necessary.

2.0 Add Loctite

3.0 With a multimeter, test the continuity between the back case, front case and the top panel. If the meter can be made to ring through, then the continuity is good.

4.0 [image: image14.jpg]


Add the DPM surround,
The surround will require  50% of the rear removing to accommodate the sensor cable.
Depending on how the cable lies one of the locating lugs may also require removing


5.0 Take a DPM window and peel off any stickers on the front, but not the protective film. Peel off the card on the back of the window to expose the adhesive, peel off the protective film that is on the DPM itself and stick the window to the DPM being careful not to trap any dust between the two.

6.0 Check the cable is superglued into the grommet

7.0 Partly close the case to see if the window looks level, this is best done with the instrument on so as to see if the display looks correct. Adjust the window and retry if necessary.

8.0 Take a case window and insert it firmly into the front case, this is best done by using the plastic bag  to avoid creating fingerprints.

9.0 Gently run a soldering iron around the edges of the window to fix it into place.

10.0 Take the front case and hold it over the top of the back, then run a line of silicone sealant down the groove on the left hand side of the back case. Gently lift the front case back over to the left hand side, trying not to touch the line of sealant just laid.

11.0 Lay sealant along the edge of the top panel, the bottom and right hand edges of the back case.

12.0 Do not forget the battery wire access in the top cover

13.0  Peel the protective label off the DPM and quickly close the case, ensuring not to trap any wires, particularly at the bottom edge. If wires do get trapped, they can be pushed in with a thin screwdriver.

14.0 Insert a case screw and tighten. Do not over-tighten.

15.0 Remove any surplus silicone 

16.0 Place the top of the instrument into two plastic clamps. They should be able to grip the instrument about 1cm from the top. Ensure they do not touch the case window

17.0 Make one final check that the edges of the case are snapped closed as far as they can go before leaving the sealant to cure, this is best done by leaving over-night. Wipe off all excess sealent before leaving including inside the battery door and on the cable.

“L”   Finishing off the Analyser


1.0 Retest the Analyser 

2.0 No sensor  0.00

3.0 With sensor reads and calibration adjusts

4.0 Remove all the excess silicone rubber from the case, including inside the battery compartment. Any rubber in the side grooves of the casing can be removed with a scalpel

5.0 Stick a back label onto the case. Stick a label on the front at the bottom. These two labels are the only two labels that will be different if the monitors have been made with customer specific labelling.

6.0 Attach a serial number label on the inside of the battery compartment over the production label. Place the battery into the compartment and close the door. 

7.0 Ensure the cable is clean and free from rubber. Clean the entire instrument with silicone polish using a soft cloth or kitchen roll.

8.0 Place the Analyser into a bag, making sure that the coiled cable is down the side of the instrument, and not across the labels, as this can leave a lasting mark on them.

9.0 Label the bag with a serial number label. Batteries should be removed prior to despatch or storage. But are best left on for a couple of days to ensure no problems develop  e.g. Drifting or draining.

10.0 None auto switch off can be enabled by connecting together both points of C2 and a green on/off switch being used.

Tools


VN202 assbly  100617.doc    6/17/2010        Page 7 of 7

