[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       



T ADAPTER – CE FILE 

Risk Assessment iaw EN ISO 14971:2000 Annex D : Possible hazards with medical devices.

	Ref.
	Hazard.

	Related part / Component posing risk.
	Sev

of

Haz.
	Like

of

Haz.
	Risk.
	Solution.
	Document referenced.
	Sev

of

Haz.
	Like

0f

Haz.
	Risk.

	D.2
	Energy hazards and contributory factors

	D.2.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.2
	Heat
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.3
	Mechanical force
	Adapter
	1
	1
	1
	Recommend not to use force when making connections
	User Instructions
	1
	1
	1

	D.2.4
	Ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.5
	Non ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.6
	Moving parts
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.7
	Unintended motion
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.8
	Suspended masses
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.9
	Patient support failure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.10
	Pressure (vessel rupture)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.11
	Acoustic pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.12
	Vibration
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.13
	Magnetic fields 

(eg. MRI)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3
	Biological hazards and contributory factors

	D.3.1
	Bio-contamination
	Adapter
	1
	2
	2
	Easy to Clean
	Cleaning Instructions
	1
	1
	1

	D.3.2
	Bio-incompatibility
	Adapter
	1
	2
	2
	Plastic Product
	Manufacturers Data
	1
	1
	1

	D.3.3
	Incorrect formulation (chemical composition)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.4
	Toxicity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.5
	Allergenicity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.6
	Mutagenicity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.7
	Oncogenicity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.8
	Carcinogenicity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.9
	Re and/or cross infection
	Adapter
	1
	1
	1
	Easy to Clean
	Cleaning Instructions
	1
	1
	1

	D.3.10
	Pyrogenicity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.11
	Inability to maintain hygienic standards
	Adapter
	2
	2
	4
	Easy to Clean
	Cleaning Instructions
	2
	1
	2

	D.3.12
	Degradation
	Adapter
	1
	1
	1
	Data Sheets
	Life of Material
	1
	1
	1

	D.4
	Environmental hazards and contributory factors

	D.4.1
	Electromagnetic fields
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.2
	Susceptibility to electromagnetic interference
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.3
	Emissions of electromagnetic interference 
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.4
	Inadequate supply of power 
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.5
	Inadequate supply of coolant
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.6
	Storage / operation outside prescribed environmental conditions
	Adapter
	1
	2
	2
	Environmental conditions as per hospital protocols
	User Instructions
	1
	1
	1

	D.4.7
	Incompatibility with other devices with which the product is intended to be used
	Adapter
	1
	2
	2
	Data Sheets
	Product by Size and suitability
	1
	1
	1

	D.4.8
	Accidental mechanical damage
	Adapter
	1
	1
	1
	Relatively robust material, if damaged, user to assess damage before use
	User Instructions
	1
	1
	1

	D.4.9
	Contamination due to waste products and/or device disposal
	Adapter
	1
	1
	1
	No special disposal required
	Risk Report
	1
	1
	1

	D.5
	Hazards resulting from incorrect output of energy and substances

	D.5.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.2
	Radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.3
	Volume
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.4
	Pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.5
	Supply of medical gases
	Adapter
	1
	2
	2
	Hospital Protocols
	User Instructions
	1
	1
	1

	D.5.6
	Supply of anaesthetic agents
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6
	Hazards related to the use of the medical device and contributory factors

	D.6.1
	Inadequate labelling
	Adapter
	1
	1
	1
	No Label required
	User Instructions
	1
	1
	1

	D.6.2
	Inadequate operating instructions
	Adapter
	1
	1
	1
	Easy to Use
	User Instructions
	1
	1
	1

	D.6.3
	Inadequate specification of accessories
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.4
	Inadequate specification of pre-use checks
	Adapter
	1
	1
	1
	Easy to Use
	User Instructions
	1
	1
	1

	D.6.5
	Over-complicated operating instructions
	Adapter
	1
	1
	1
	Easy to Use
	User Instructions
	1
	1
	1

	D.6.6
	Inadequate specification of service and maintenance
	Adapter
	1
	1
	1
	Easy to Use
	User Instructions
	1
	1
	1

	D.6.7
	Use by unskilled / untrained personnel
	Adapter
	1
	2
	2
	Easy to Use
	User Instructions
	1
	1
	1

	D.6.8
	Reasonable foreseeable misuse
	Adapter
	1
	2
	2
	Easy to Use
	User Instructions
	1
	1
	1

	D.6.9
	Insufficient warning of side effects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.10
	Inadequate warnings of hazards likely with re-use of single use devices
	Adapter
	2
	2
	4
	Product Data Sheets
	User Instructions / Cleaning Instructions
	2
	1
	2

	D.6.11
	Incorrect measurement and other metrological aspects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.12
	Misrepresentation of results
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.13
	Incompatibility with consumables / accessories / other devices
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.14
	Sharp edges or points
	Adapter
	2
	2
	4
	If damaged, user to assess damage before use
	User Instructions
	2
	1
	2

	D.7
	Inappropriate, inadequate or overcomplicated user interface (man/machine communication)

	D.7.1
	Mistakes & judgement errors
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.2
	Lapses and cognitive recall errors
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.3
	Slips & blunders (mental or physical)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.4
	Violation or abbreviation of instructions, procedures etc
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.5
	Complex or confusing control system
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.6
	Ambiguous or unclear device state
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.7
	Ambiguous or unclear presentation of settings, measurement, or other information
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.8
	Misrepresentation of results
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.9
	Insufficient visibility, audibility or tactility
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.10
	Poor mapping of controls to action or of displayed information to actual state
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.11
	Controversial modes or mappings as compared to existing equipment
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8
	Hazards arising from function failure, maintenance and ageing and contributory factors

	D.8.1
	Erroneous data transfer
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.2
	Lack of, or inadequate specification for maintenance including post maintenance functional tests
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.4
	Inadequate maintenance
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.5
	Lack of adequate determination of end of device life
	Adapter
	1
	1
	1
	User decision based on clarity of plastic & tolerable damage
	Risk Analysis Report
	1
	1
	1

	D.8.6
	Loss of electrical integrity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.7
	Loss of mechanical integrity
	Adapter
	1
	1
	1
	User decision based on clarity of plastic & tolerable damage
	Risk Analysis Report
	1
	1
	1

	D.8.8
	Inadequate packaging (contamination and / or deterioration of the device)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.9
	Re-use and/or improper re-use
	Adapter
	2
	2
	4
	Product Data Sheets
	User Instructions
	2
	1
	2

	D.8.10
	Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
	N/A
	1
	1
	1
	
	
	1
	1
	1


E. Risk Management BS14971 Annex D_T Adapter_Reusable & Disposable
14/02/2005
    Page 6
Severity of hazard: 1 = Insignificant, 2 =  Tolerable, 3  = Critical, 4  =  Intolerable.  Probability of event: 1 = Improbable, 2 = Occasional, 3 = Likely, 4 = Highly likely.  

Risk calculated as severity of hazard x probability of event, 1 – 16.  Further explanation of risk management policy – see Risk Management Policy & Definitions.


