	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Oxygen Sensor manufacture

	R-22OMS sensor

VM3/COP/40.16

	Date: 30-Sep-03
	Revision Date: 2-Jul-10
	Issue 1

	Only to be attempted by trained personel



	Parts Required

	Quantity
	Description
	Part No.
	

	1
	R-22SHO
	
	

	2
	Serial number Labels
	
	

	1
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	


Warning  Use Soldering Iron set at 300C Do Not overheat the PCB on the Sensor

Procedure

1. Remove the R22SHO from the Gas barrier bag, 

2. Test the sensor 25mV +/- 30% 

3. Using procedure VM3/COP40.11 “Modification of sensor case. Remove excess case

4. Break the seal of the conformal coating with a scalpel and remove the shroud with a pair of pliers. Do not dislodge the PCB during this action

5. Remove the white surround

6. Bend the pins ( there will be no black holder around the pins

[image: image1.jpg]



7. Leave at an angle of approximately 45degrees to the PCB to allow the OMS connector to fit

8. Check the output 25mV ( Can be checked on an OMS profession Dive Oxygen Analyser 

9. Remove the Teledyne Label

10. Add the Vandagraph label to the sensor and the container (Now White incorporating printed serial number Date code and bar code
[image: image2.jpg]@ VANDAGRAPH LTD cuassrzzous TN
M ANDACRASL TP Crae
Vs ane
T
UKL S

WARNING! CONTAINS CAUSTIC! CAUSES SEVERE BURNS! MAY BE FATAL & SWALLOWEL.







11. Place in Gas barrier bag and label

6739%5B1%5D.doc       02/07/2010      Page 1 of 1

