[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




[image: image2.png]




T ADAPTER – CE FILE 

Risk Management Policy & Definitions.

Viamed Policy on Risk Management.
Annex I, paras 1, 2 & 6 of the MDD/93/42 address risk management.

	1. Definitions of hazard severity and  event probability are :
	
	Rating.
	Hazard severity.
	Event probability.

	
	
	1
	Insignificant.
	Improbable.

	
	
	2
	Tolerable.
	Occasional.

	
	
	3
	Critical.
	Likely.

	
	
	4
	Intolerable.
	Highly likely.


2. Solutions are implemented to reduce hazard severity from level 4 (intolerable) identified at the conceptual stage to level 2 (tolerable) or lower.  

If hazard severity cannot be reduced from level 4 (intolerable), the product is abandoned.

3. Hazard severity and event probability which cannot be reduced from level 3 (critical) and level 3 (likely) / level 4 (highly likely) respectively are reviewed by the MD and either:

i. Signed off if considered acceptable when weighed against the benefits.

ii. The product abandoned if not.

4. Effectiveness of implemented risk management solutions is implied by a reduction in ratings.

5. The overall risk rating formed by the product of hazard severity x event probability cannot alone be the deciding factor of whether product development continues or is abandoned; a low risk occurring regularly would be considered less serious than a high risk occurring occasionally.

6. See below for action taken for hazard severity x event probability combinations:

	Hazard severity
	Event probability
	Risk 

product
	Level

of risk
	Action taken

	1
	1
	1
	Low

[image: image1.jpg]
High
	No further action.

	1
	2
	2
	
	No further action.

	2
	1
	2
	
	No further action.

	2
	2
	4
	
	No further action.

	1
	3
	3
	
	Reduce event probability or MD review of benefit to risk.

	3
	1
	3
	
	Reduce hazard severity or MD review of benefit to risk.

	1
	4
	4
	
	Reduce event probability or MD review of benefit to risk.

	2
	3
	6
	
	Reduce event probability or MD review of benefit to risk.

	3
	2
	6
	
	Reduce hazard severity or MD review of benefit to risk.

	2
	4
	8
	
	Reduce event probability or MD review of benefit to risk.

	3
	3
	9
	
	Reduce hazard severity or MD review of benefit to risk.

Reduce event probability or MD review of benefit to risk.

	3
	4
	12
	
	Reduce hazard severity or MD review of benefit to risk.

Reduce event probability or MD review of benefit to risk.

	4
	1
	4
	Intolerable
	Reduce hazard severity or abandon the product.

	4
	2
	8
	Intolerable
	Reduce hazard severity or abandon the product.

	4
	3
	12
	Intolerable
	Reduce hazard severity or abandon the product.

Reduce event probability or MD review of benefit to risk.

	4
	4
	16
	Intolerable
	Reduce severity of hazard or abandon the product.

Reduce event probability or MD review of benefit to risk.


7.   Risk analysis / management is reviewed periodically to ensure safety of products & users.
E. Risk Analysis Definitions_T Adapter
11/02/2005                                                                                                                                                      Page 1

[image: image2.png]