	VANDAGRAPH
 SENSOR TECHNOLOGIES

Company OPERATING PROCEDURE

	Oxygen Sensor manufacture

	R-22AEX   Expedition Sensor

VM3/COP/40.47

	Date: 30-Sep-03
	Revision Date: 29-Jul-10
	Issue 1

	Only to be attempted by trained personel

	Used in Oxycheq Expedition  Monitor    Pt Number OA-3 440006



	[image: image1.jpg]


[image: image2.jpg]




	Parts Required

	Quantity
	Description
	Part No.
	

	1
	R22  Sensor ( R-22MED: R-22A: R-22BUD are all equivalent)  

R-22A 1st choice
	
	

	2
	Pins Farnell 143-116
	9071046
	

	1
	Shell 3pin    Farnell 143-093
	9071033
	

	1
	Red wire 5 cm  or Red
	9070130
	

	1
	Black wire 5 cm
	9070131
	

	1
	Pre-assy ( above 4 components)
	9070110
	

	2
	Serial number Labels
	
	

	1
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	

	1
	Gas Barrier bag
	
	


Warning  Use Soldering Iron set at 300C Do Not overheat the PCB on the Sensor

Method.

[image: image3.jpg]



1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Remove excess outer case unto level of PCB. VM3COP40.11

4. Remove the Molex and connect wires

5. Connect Black & White wires

6. Pay attention to wire positions on the PCB Black to BLK: White to RED

7. The lock position on the molex

8. The center is left unconnected

9. Flat side of molex  White to the  LEFT

10. Label sensor with R-22AEX

11. Wait until sensor cools to room temperature

12. Test for output 7-13mV.

13. Can be tested in Expedition analyser

14. Add R-22AEX label to sensor,& bag

15. Reseal in serial numbered gas barrier bag

16. Ensure the O2 sensor is booked out in the stock books.

(*).  If the O2 sensor fails the output specification test, it should be replaced in it’s original packet and re-boxed.  It should then be placed in a docket and given a Repair sheet, tested and returned to supplier under the correct procedures.

r-22AEX.doc         29/07/2010                       Page 1 of 2

