
Review of the VN202 LITE Enclosure

The overall appearance is good

The look and feel are good and it feels balanced

The six retaining screws make the instrument feel solidly assembled

Why are there eight small holes in the space between the batteries. Is this for a future alarm and if so how will this effect the waterproofing

Can you confirm that we have a seal all around the case.

Can you confirm that we have a seal all around the battery cover

Due to the experience with the Tek Ox. Are the screw hole mounts substantial.

I appreciate they are not going to be removed as frequently.

Can we use the same screws as the Tek ox 5mm:7.5mm:10mm to make service easier

There is no waterproofing around the jack socket. I would feel happier if this could be enclosed with the wires siliconed around an exit hole. 

This instrument will be subjected to sea water and sea air. In the tropics this will cause corrosion on the socket within 6 months if the mail connector is not in place and screwed tight.

I would like to see how you are going to connect the PCB to the rest of the instrument. 

I have added a Tek-Ox PCB to check the aesthetics with the display.

The cables from the pcb connector have to travel through 180 degrees. 


