P	Cp	Sc	Statement on report
6	1		There were no specific re-breather diving sensors as no one manufactured re-breathers for the recreational diving market’ .Until 1993 only military and Technical divers were involved usually with ex military re-breathers.
6			The R22 MED ( formerly R-22) was supplied to APV for evaluation / R-22A is an Automotive sensor and should not have been supplied to a diving company. / These sensors were not supplied by Vandagraph
6			Quality / The R-22MED was developed with APV input into the R-22BUD / Non –conformal coating / Removal of grill / Hydrophobic membrane / Vent holes increased to 2mm / Output of the sensor changed from 7-13mV with a nominal 10.5 mV to 8mV-13mV .The sensor was trimmed to fit the APV software. And included the APV safety factors. / These changes were possible because of the constant monitoring and feedback from divers using the APV Inspiration
6			QA route / Teledyne QA / Vandagraph QA – output in air & visual checks / APV QA / Quality statement is not supported by Vandagraph records
6			Teledyne introduced the R-22D / It was in response to other companies requesting sensors for re-breathers. / It was not identical to the R-22BUD. / It had a different output range 7mV-13mV R-22BUD had 8mV-13mV. / APV had already carried out temperature checks in the Inspiration using the original temperature compensation circuit and had designed the software to enable good tracking. APV were upset when Teledyne used all there input to release a general R-22D.
6			Competition supply / As with any manufacturer APV approached Teledyne to manufacture a specific APV OEM sensor the R22-DAP to specifically match their needs. / Teledyne had to sign a NDA and give an undertaking only to manufacture this sensor for APV. ( In my presence TC, KK SB) Although this initially may have been intended for a chip sensor after the singing the DAP version ( originally the test samples were manufactured by IT) was produced. / Teledyne & APV agreed to the technical specification and feasibility
6			Sensor life / Not borne out by Vandagraph information
6			Temperature compensation / The original Inspiration software was developed with the R-22MED. Temperature effects can only be evaluated if tested in the Inspiration.
7			Linearity variations not observed by Vandagraph until the sensor is nearly exhausted. Sensors tested by Vandagraph can only be tested at 21% and 100%. Linearity is checked to 3%. It is possible to test between 21 & 100% if required / Several factors come in to play on this test notably ambient pressure and the time to clear the 100% oxygen from the sensor where it can get trapped unless a turbulent gas flow is used.
7			It is easy to manufacture small quantities of sensors with a exceptionally high level of QA / Moving production from small quantities which can be easily controlled mass production of thousands makes continuous high levels of QA harder to employ and monitor particularly in a labour intensive product. / This is reflected in the ISO standards and subsequent external Audits / When the effects of practical QA application and monitoring are understood the product specification is written to accommodate the foreseeable human variations that can occur. I.e. labour skills, and sub contractor tolerances. / In the case of APV and diving sensors the user is advised to change the sensor every 12 months with 18 months being the maximum period of use / In this period the level of diving problems experienced is minimal / Over 18 months the level of problems accelerates / By the nature of their manufacture all sensors are subject to possible problems either in individual sensors or in batches / Many factors can affect the life of a sensor. / In reality the recreational diving sensor is subjected to many environmental changes such as pressure and temperature which cause stress. / Since 1996 continuous monitoring of actual dived sensors as opposed to theoretical laboratory tests and interaction between APV Vandagraph and Teledyne has revealed many areas where the original sensor could be improved and theoretically made more reliable. Where practical, and bearing in mind the divers attitude to purchasing cheaper sensors, these improvements have been incorporated and QA has been extended to create the R-22-2BUD and the R22-DAP sensors / Since 1965 leakage in a proportion of sensors has been observed and taken into account. Due to manufacturing techniques and components used the levels of leakage in practice have substantially improved but have not been eradicated. / NB. It has been observed and reported that sensors do not last for there theoretical calculated life when diving. I have assumed this to be caused by the continuous stressful changes the sensor is exposed to. APV recommend strongly changing the sensors every year. This is not followed by all divers.
7			True R-22A is an automotive sensor
7			Response time is checked within the Inspiration and slow sensors are rejected. / Response time controlled at manufacture by choice of membrane. / Very few sensors over the last 30 years have been diagnosed a slow response. / (Figures for the medical sensors may be recoverable if necessary) / Diving sensors are frequently reported as being slow. This has in many cases been traced to moisture. 5-7 yrs ago I asked several divers reporting this on the web to return the sensor to me specifically for examination. I have had none returned.
7			Drift. Not borne out by the facts on sensors returned to Vandagraph / Sensors have been kept in sealed bags from 2002 some are still giving an output
			Shock Cannot comment on this. Fast decompression is not recommended in diving
			Chamber lock out shows a little bit of bias in the comments
8			Usually microscopic leaks prevent the sensor working and would be displayed as a failure to calibrate or unstable baseline. All sensors in all environments are expected to be calibrated every 8 hours maximum / In the past it was calculated that 50% of sensors would eventually die due to leakage. I do not know if this is still true. Most of the sensors we see are still under warranty.
8			Storage. We have sensors stored since 2000 ? still with an output in air
8			Life Tests Accelerated tests are theoretical and no substitute for long-term monitoring and data collection
			Storage : Not borne out by Vandagraph figures
8			Teledyne and Vandagraph have been consulted many times on the use of sensors and NDA have been signed and OEM sensors have been produced / It is unethical to pass on one company’s findings usually at great cost to another company. It is always easier for a second source to be cheaper in the market place as he has no follow up procedures to follow. / If this company had followed the correct procedures I would have examined there comments and where required persuaded Teledyne of the need for change.
8			Cost difference is due to the levels of QA and follow up.
			Admits the small size of sample
9			If Vandagraph was approached to manufacture a sensor we would request a detailed specification. Changes can be made to ensure ttey are different and unusable in the Inspiration. The DAP connector is freely available from electronic suppliers and could be supplied by the OEM as a parts issue.APV are only concerned that Teledyne to not offer an identical sensor to the R-22DAP which can be released uncontrolled into the market and have a price advantage without the APV responsibility. / Many web based suppliers now have very little technical knowledge on sensors particularly in the diving environment.
			Sensor temperature is effected by temperature of the electrolyte (2.5%/1C / The best location for temperature compensation is in the electrolyte. This speeds up the sensors reaction and reduces the effect of a temperature gradient between the anode and the temperature compensation circuitry.
			This is a bold statement made on the testing of 12 sensors. / There have been at least 28K diving sensors supplied by Vandagraph alone
			The internet history has to be read to be believed / One diver has a problem and 10,000 experts offer an opinion. / On the Inspiration website I or Martin let it go for a few days then jump in with a statement and it all goes quiet. / It is unbelievable how the divers are not only abusing the sensors but actually boast about it
			Vandagraph are responsible for the UK and Europe for Teledyne diving products / Vandagraph started the recreational diving side of Teledyne ( A Pitcelli was party to a lunch when I predicted the possible growth market. Seasoned Teledyne employees refuted this / No approach has been made to Vandagraph by this organisation / History shows Teledyne do respond to valid constructive criticisms and requests. Vandagraph has published evidence of how much we care for our customers.The report writer deliberately ignores Vandagraph or has no working knowledge of the marketplace.
9	3		No data on actual dived sensors
10	6		10% Humidity / If the condensate lies on the front membrane it reduces the surface area of the diffusion path. This slows down the sensor response. When the membrane is completely covered the sensor will measure the oxygen dissolved in the condensate. / Salt water if allowed to dry will: / contaminate the front membrane and slow the sensor down / short circuit the components on the temperature compensation PCB. / Any copper wire exposed will turn black and wires will become brittle. / It is recommended that any salt water be flushed out of the system as quickly as possible with fresh water and allowed to dry naturally / Calibration of the sensors should then be carried out in Air and 100%
			90C is out of specification
			CO2 reacts with the electrolyte KOH. / This is a chemical reaction in the sensor / Sensors can tolerate short-term exposure or physiological levels. / Special designs of sensor are required to tolerate higher levels of cO2
			ISO 9000:2001 and ISO13485:2003 are standard in Teledyne and Viamed/Vandagraph
11	7		Thermal conductivity and Ultrasonics have been omitted
	8		Viamed & Vandagraph have been omitted
	8+		Seems a strong bias towards AII
13	9.2	1	Has ignored the Teledyne original patents and papers describing sensors

			Organic contamination is controlled by the QA procedures 12 samples would not be statistically significant considering the numbers manufactured. Also amount of contamination may be acceptable within the life specification?
			Teledyne Cell housing has varied
14	9.3	1	Sensors also fail because they leak. This has reduced over the years but is not completely eradicated.
			Teledyne calculate shelf line differently. I believe it used to be 5% loss of useful life for every year in the gas barrier bag. ( notes on Life & warranty)
	9.3	2	b. Failure of the temperature compensation network is rare
14	9.3	3	Sensor membrane; only damaged if the sensor is dived outside of its specification or the recommended diving envelope. I have used this in the past to describe that sensor can get bent just like a dive and so should always be dived correctly
15	9.4	1	Too small a sample to be statistically relevant / One problem not seen by Deas is the failure on packing to completely seal the gas barrier bag. This can happen unknown to the packer due to wear on the heat sealer. Our second line QA prevents this in most cases
	9.4	2	This is not entirely borne out in the results we have on returned sensors / I also believe this was reduced with the introduction of the Rhodium cathode
		4	This only happens if the diver dives outside the recommended diving envelope
19	9.4.	5	This was observed several years ago and was the result of too much non-conformal coating ( human error) Visual checks were instigated to eradicate this problem. Secondly the ventilation holes were increased in diameter. Visual ckeks in Vandagraph and APV are still in place. This problem was restricted to a small number of sensors and was only discovered due to the close relationship of APV & Vandagraph with actual divers who could be trusted to give the true facts. / The key to this discovery was a respected dive decompressing on a fixed surface buoy and observing the PPO2 readings increase as he moved up and down due to wave motion .Subsequent investigation of the spikes observed by the diver were traced by APV and Teledyne to blocked holes. / This would probably never have been discovered without the interaction. / It was also rare because the sensors have two holes and in many sensors the the gas is able to equalise around the edges of the PCB
16	9.4	6	It is not safe to consider any comments by divers on the Internet. They are all instant experts. There are a few well trained and experienced divers and these have been used to evaluate the more spurious comments that have appeared.
		7	Teledyne or Vandagraph or APV may be able to produce failure figures as the results are obvious due to the KOH / This statement is misleading as a massive KOH leak prevents the sensor functioning.
		8	In the Inspiration this type of fault would show BEFORE the unit was dived. / This type of fault has to be rare and a probable batch problem or human mix up. I also believe it would show up on the Teledyne computerised tests / It could refer to the R-22A
16	10		Need to establish this point as I not sure this statement is correct. The RoHS compliance regulations are continually under review and in many areas we cannot get definitive answers from our UK Dept. of Trade and Industry.
17	11		Would not like to comment on this section
20	12		Not in a position to comment on this section
	13		I need to discuss these tests with Teledyne and there relevance to actual re-breather dividing with APV. APV have many tests rigs and have acquired a CE mark for their re-breathers. They also work closely with the Royal Navy. / They also have a very close relationship with the UK Health & Safety Executive and I believe you will find APV have implemented any HSE recommendations.
	13.1		Not sure this is a valid remark. The molex connectors were used for several years. Who tested the Inspiration after the fatal accident. APV state that no Teledyne sensor has been found to be the cause of a fatal accident to date / NB the current SMB connector is not 100% . WE have certain reservations about the consistency and QA on this connector. The R-17D has proved to be successful and gained another re-breather the CE mark. The R-17MED was recommended to APV originally but space was a problem. They were also concerned that Jack plugs can exhibit a dead spot. This seems to have been overcome with the second re-breather manufacturer. Unfortunately I cannot divulge this to others. I can only recommend the use of R-17D sensors rather than R-22D's
	13.2		This figure needs confirmation. I doubt its veracity. I would think other components were more at risk at these temperatures
	13.3		This test is valid as an extreme but his backup notes are questionable. / Most equipment on RIBS is tied down to prevent bounce and the equipment being bounced overboard. Most RIBs have a maximum speed below 40mph. / The greatest drop is from a dive boat (converted fishing boat) Here the height can be more than 3 mtrs but the divers body takes most of the shock.
			It is recommended that diving sensors be changed every 12 months . / It is a hostile environment and although the system is sealed the diver breaks this seal when he prepares the re-breather. Opening the re-breather in Sea air allows air contamination. Leaving for 6 months only allows the corrosion time to take effect.
28	13.4		APV meets all relevant standards
	13.5		Sensors should be used at 1Bar (1 atmosphere) so it would be more logical to accelerate the tests using 100% oxygen.
29	14		This is only valid in Air. In 100% oxygen the sensor is restricted to 5 Bar. It flattens out at 350mV output / The front and rear membranes should always be at the same pressure
29	14.1		R17 and R22-A are not diving sensors
	14.2.1		As I see at least 30K sensors per year pass through my companies destined for Automotive, Medical, and diving markets. My experience does not confirm this statement as being reliable. Much depends on how the sensors were stored. / How is failure rate calculated at 100hrs / Any sensor surviving until lead exhaustion will show a ceiling effect as the internal impedance increases the output will flatten out and cease. / What is significant drift?
30	14.2.2	1	This has changed
		2	Over all sensors no shelf life problems have been reported by users
		3	Wide spread reports but by who. Divers? Gain access to the forums and the level of knowledge on sensors will quickly be revealed. There are a lot of so called experts but the information is riddled with conjecture and a lack of actual knowledge. Vandagraph has always offered a 12 month no quibble warranty on all diving sensors. If they failed under warranty and were returned examination to the cause could be established. The reference provided exposes an author with limited knowledge of sensors and there diving application
	Fig 8.1		Has this sensor been evaluated by Teledyne.
32	Fig 8.4		Labelling is historic. It can be changed to meet any ISO or CE standard. The customer controls the label information
	14.2.3		Many definitive conclusions from a small sample of old sensors.
	14.3		R22A is an automotive sensor. / The distributor was not Vandagraph or Viamed.
40			Did Teledyne carry out RMA44213
			The Teledyne report shows that Teledyne have responded to finding a problem
41	14.3.7	1	This statement is confusing. I assume they mean they re-tested an R22A. / The distributor is not in a position to make these statements unless confirmed by Teledyne.
42			APV QA the sensors in an actual re-breather test rig.
50	14.4.7	1	There is no R17DA sensor
			A more likely reason for the high output is the increase in internal pressure
	14.5		The R22-DAP is an OEM sensor and cannot be made available to other users. / This applies across the range e.g R22CC manufactured for Critikon, R17VAN manufactured for Vandagraph.
55	15.3		He quotes ceiling faults as the correct failure mode