[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       Apgar Timer II

– CE FILE            DATE:

Risk Assessment iaw EN ISO 14971:2000 Annex D : Possible hazards with medical devices.Ammended to match EN14971 :2009


	Ref.
	Hazard.


	Related part / Component posing risk.
	Sev

of

Haz.
	Like

of

Haz.
	Risk.
	Solution.
	Document referenced.
	Sev

of

Haz.
	Like

0f

Haz.
	Risk.

	D.2
	Energy hazards and contributory factors

	D.2.1
	Electricity
	None
	1
	1
	1
	Non required
	ISO9000:2000

ISO14971:2009

ISO13471:2009

Y1 Design History

YZ Design sFiles
	1
	1
	1

	
	Line voltage
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Leakage enclosure
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Leakage earth
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Patient Leakage
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Electric Fields
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.2.2
	Heat
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	High Temperature
	None
	1
	1
	1
	Non required
	IEC601/1  PtII
	1
	1
	1

	
	Low Temperature
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.2.3
	Mechanical force
	Instrument
	2
	1
	2
	May be dropped by user
	F User Manual
	1
	1
	1

	
	Gravity : Falling
	Instrument
	1
	1
	2
	Non required
	
	1
	1
	1

	
	Suspended masses
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Stored energy
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Torsion,Shear & Tensile
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	High Pressure  Fluid injection
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.2.4
	Ionising radiation
	None
	1
	1
	1
	Non required
	Z12 No Ionisation 
	1
	1
	1

	D.2.5
	Non ionising radiation
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.2.6
	Moving parts
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	
	Moving & positioning patient
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	D.2.7
	Unintended motion
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	D.2.8
	Suspended masses
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	D.2.9
	Patient support failure
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	D.2.10
	Pressure (vessel rupture)
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	D.2.11
	Acoustic pressure
	None
	1
	1
	1
	Non required
	D2 Device description
	1
	1
	1

	
	Ultrasonic energy
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Infrasound energy
	None
	1
	1
	1
	Non required
	D2 Device descriptionD2 
	1
	1
	1

	
	Sound
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.2.12
	Vibration
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.2.13
	Magnetic fields 

(eg. MRI)
	None
	1
	1
	1
	Do not use in MRI
	D2 Device description

F5 User manual
	1
	1
	1

	D.3
	None

	D.3.1
	Bio-contamination
	None
	2
	2
	4
	Easy to clean 
	Z6 Bio-compatibility

F5 Instruction manual
	1
	1
	1

	
	Bacteria
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Viruses
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Other agents (prions)
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.2
	Bio-incompatibility
	c
	1
	1
	1
	No adverse matrials used
	Z6 Bio-compatibility
	1
	1
	1

	D.3.3
	Incorrect formulation (chemical composition)
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.4
	Toxicity
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.5
	Allergenicity/ irritancy
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.6
	Mutagenicity
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.7
	Oncogenicity
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.8
	Carcinogenicity
	None
	1
	1
	1
	Non required
	Z11 Carcinogenic substances
	1
	1
	1

	D.3.9
	Re and/or cross infection
	Instrument & Leads
	1
	1
	1
	Non required
	O1 Sterilisation

F6 Cleaning instructions

F5 Instructions for use 
	1
	1
	1

	D.3.10
	Pyrogenicity
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.3.11
	Inability to maintain hygienic standards
	Instrument & Leads
	2
	2
	4
	Easy to clean – cleaning instructions in user manual
	O1 Sterilisation

F6 Cleaning instructions

F5 Instructions for use 
	1
	1
	1

	D.3.12
	Degradation
	Instrument 
	2
	1
	2
	Care instructions given in the user manual
	(F) User manual 
	1
	1
	1

	
	Chemical
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Acids or Alkalis
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Contaminates
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Processing aids
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Cleaning, disinfection
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Testing aids
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Medical gases
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Anaesthetic products
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.4
	None

	D.4.1
	Electromagnetic fields
	None
	1
	1
	1
	Instrument Design
	
	1
	1
	1

	D.4.2
	Susceptibility to electromagnetic interference
	None
	1
	1
	1
	Instrument Design
	D4 EMC
	1
	1
	1

	D.4.3
	Emissions of electromagnetic interference 
	None
	1
	1
	1
	Instrument Design
	D4 EMC
	1
	1
	1

	D.4.4
	Inadequate supply of power 
	None
	1
	1
	1
	Battery
	
	1
	1
	1

	D.4.5
	Inadequate supply of coolant
	None
	1
	1
	1
	No coolant
	
	1
	1
	1

	D.4.6
	Storage / operation outside prescribed environmental conditions
	Instrument
	2
	1
	2
	No specific environmental storage / operating conditions imposed
	
	1
	1
	1

	D.4.7
	Incompatibility with other devices with which the product is intended to be used
	None
	1
	1
	1
	Non required Battery : Stand alone 
	
	1
	1
	1

	D.4.8
	Accidental mechanical damage
	Instrument & Leads
	2
	1
	2
	May be damaged if dropped”
	F5 User manual
	1
	1
	1

	D.4.9
	Contamination due to waste products and/or device disposal
	None
	1
	1
	1
	Non required
	Electronic waste disposal required. RoHS

F6 labels
	1
	1
	1

	D.5
	None

	D.5.1
	Electricity
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.5.2
	Radiation
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.5.3
	Volume
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.5.4
	Pressure
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.5.5
	Supply of medical gases
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.5.6
	Supply of anaesthetic agents
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.6
	Hazards related to the use of the medical device and contributory factors

	D.6.1
	Inadequate labelling
	Label
	2
	1
	2
	No instruction label on instrument – Instructions for use attached
	F7 Labels
	1
	1
	1

	D.6.2
	Inadequate operating instructions
	User Instructions
	2
	1
	2
	Instructions for use comprehensively cover the product.  Instrument only for use by trained and qualified personnel. 
	F5 User Instructions
	1
	1
	1

	
	Inadequate description of performance
	None
	1
	1
	1
	Non required
	M3 SpecificationF5 User Instructions
	1
	1
	1

	
	Inadequate specification of intended use 
	None
	1
	1
	1
	Non required
	F5 User Instructions
	1
	1
	1

	
	Inadequate disclosure of limitations
	None
	1
	1
	1
	Non required
	F5 User Instructions
	1
	1
	1

	D.6.3
	Inadequate specification of accessories
	None
	1
	1
	1
	Non required
	F5 User Instructions
	1
	1
	1

	D.6.4
	Inadequate specification of pre-use checks
	Instructions for Use
	2
	2
	4
	Instrument must be pre-use checked – Instructions for Use refer. 
	F5 User Instructions

F9 Training
	2
	1
	2

	D.6.5
	Over-complicated operating instructions
	Instructions for Use
	1
	1
	1
	Non required
	F5 User Instructions
	1
	1
	1

	D.6.6
	Inadequate specification of service and maintenance
	Instrument


	2
	1
	2
	No service or maintenance except batteries
	G2 Service Manual

G1 maintenance

G3 Service Procedures
	1
	1
	1

	D.6.7
	Use by unskilled / untrained personnel
	Instrument
	2
	2
	4
	Instructions for use comprehensively cover the product.  Instrument only for use by trained and qualified personnel. 
	F5 User Instructions

F9 Training
	2
	1
	2

	D.6.8
	Reasonable foreseeable misuse
	Instrument
	1
	1
	1
	Batteries removed for long periods of non-use
	
	2
	1
	2

	D.6.9
	Insufficient warning of side effects
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.6.10
	Inadequate warnings of hazards likely with re-use of single use devices
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.6.11
	Incorrect measurement and other metrological aspects
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.6.12
	Misrepresentation of results
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.6.13
	Incompatibility with consumables / accessories / other devices
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.6.14
	Sharp edges or points
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.7
	None

	D.7.1
	Mistakes & judgement errors
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Pre-use check always before use.
	
	1
	1
	1

	
	Incorrect or inappropriate output or functionality
	None
	1
	1
	1
	.
	
	1
	1
	1

	
	Erroneous data transfer 
	None
	1
	1
	1
	 check always before use.
	
	1
	1
	1

	
	Loss or deterioration in function
	None
	3
	1
	3
	Non required
	
	1
	1
	1

	D.7.2
	Lapses and cognitive recall errors
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Pre-use check always before use.
	
	1
	1
	1

	D.7.3
	Slips & blunders (mental or physical)
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Pre-use check always before use.
	
	1
	1
	1

	
	Rule based failure
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Knowledge based failure
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	
	Routine violation
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.7.4
	Violation or abbreviation of instructions, procedures etc
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Pre-use check always before use.
	
	1
	1
	1

	D.7.5
	Complex or confusing control system
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.7.6
	Ambiguous or unclear device state
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Pre-use check always if device state unclear.
	
	1
	1
	1

	D.7.7
	Ambiguous or unclear presentation of settings, measurement, or other information
	None
	1
	1
	1
	No complex user onterface
	International symbols
	1
	1
	1

	D.7.8
	Misrepresentation of results
	None
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual.
	
	1
	1
	1

	D.7.9
	Insufficient visibility, audibility or tactility
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.7.10
	Poor mapping of controls to action or of displayed information to actual state
	None
	1
	1
	1
	Only for use by qualified & trained personnel – warnings in user / technical manual.
	
	1
	1
	1

	D.7.11
	Controversial modes or mappings as compared to existing equipment
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.8
	None

	D.8.1
	Erroneous data transfer
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.8.2
	Lack of, or inadequate specification for maintenance including post maintenance functional tests
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.8.4
	Inadequate maintenance
	Instrument
	1
	1
	1
	Instrument will not pass pre-use check if servicing required.
	(F) User manual 
	1
	1
	1

	D.8.5
	Lack of adequate determination of end of device life
	Instrument
	1
	1
	1
	User decision based on damage or cost effectiveness of repair. Instrument will not pass pre use check if badly damaged or requiring servicing.
	(F) User manual 
	1
	1
	1

	D.8.6
	Loss of electrical integrity
	None
	1
	1
	1
	Non required
	
	1
	1
	1

	D.8.7
	Loss of mechanical integrity
	Instrument
	1
	1
	1
	User decision based on damage or cost effectiveness of repair.  Instrument will not pass pre-use check if badly damaged.
	
	1
	1
	1

	D.8.8
	Inadequate packaging (contamination and / or deterioration of the device)
	Instrument 
	1
	1
	1
	Bubble wrapped and boxed prior to despatch.
	(M) Packaging


	1
	1
	1

	D.8.9
	Re-use and/or improper re-use
	Instrument 
	1
	1
	1
	User decision on suitability for next use based on pre use check & tolerable mechanical damage.
	
	1
	1
	1

	D.8.10
	Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
	None
	1
	1
	1
	Non required
	
	1
	1
	1


Annex D updated Apgar Timer II.doc

11/01/2011                          Page 10 of 10
Severity of hazard: 1 = Insignificant, 2 =  Tolerable, 3  = Critical, 4  =  Intolerable.  Probability of event: 1 = Improbable, 2 = Occasional, 3 = Likely, 4 = Highly likely.  

Risk calculated as severity of hazard x probability of event, 1 – 16.  Further explanation of risk management policy – see Risk Management Policy & Definitions.


