	     [image: image2.emf]
	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire  BD20 7DT - United Kingdom
Tel: +44 (0)1535 634542     Fax: +44 (0)1535 635582

Email: info@viamed.co.uk   Website: www.viamed.co.uk



Microstim DB3 Supramaximal Nerve Stimulator
Part number: 2510001

[image: image1.png]


            

               Caution: 

This is a non-standard version unit and is easily identified as it has green switches; the standard version (part number 2510000) has red switches.
Functionality of the 2510001 stimulation is as follows:

T.O.F. Now latching switch – Continuous cyclic output, T.O.F. which repeats after a pause of 8 seconds - for a total period of 5 minutes.

D.B.S. Momentary switch - Functionality as with standard DB3

P.T.C.  Momentary switch - Functionality as with standard DB3

1 Hz . Now latching switch – Continuous output repeats for a total period of 5 minutes.

NOTE: With either of the revised T.O.F. and 1Hz functions the latching switches need to be reset to the middle position otherwise another function cannot be selected.

After use - if either of the T.O.F and 1 Hz functions are left in the latched (On) position there is still a small current drain, even after the five minute time-out period. This current consumption will drain the battery unless the switches are returned to the middle (Off) position.

19th January 2010
[image: image2.emf]