	Product Modification - Labelling




Part number: 0310030
Model: Tom Thumb TT480

Problems with the current laminated paper version ‘CAUTION’ label:

a) Does not look professional and is easily damaged physically and if moisture seeps through the edge of the cut lamination.

b) Time consuming and expensive to manufacture.

c) Vertical positioning only applicable to one customer.

d) Possible sharp lamination edges. 

[image: image1.jpg]CAUTION

( \, Must be used with a flow

| j/ meter. DO NOT exceed
15 litres/minute.

Read instructions

prior to use.





Page 1 of 2
New version label:

a) Easy to apply and cosmetically looks in keeping with the product.

b) Easy to clean, not easily damaged by moisture.

c) Low cost – no labour costs.

d) Positioning more appropriate during actual usage.

e) Litres per minute quantity increased from 10 to 15 to be in line with the operating instructions.

[image: image2.jpg]



New labelling (part number 0390018) implemented as of 14/02/11.

Label produced by Pendle Signs & Plastics
S Nixon

14/02/11

Page 2 of 2

