[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       

                                             CE file : Headboxes.

Risk Assessment iaw EN ISO 14971:2000 Annex D : Possible hazards with medical devices.

	Ref.
	Hazard.

	Related part / Component posing risk.
	Sev

of

Haz.
	Like

of

Haz.
	Risk.
	Solution.
	Document referenced.
	Sev

of

Haz.
	Like

0f

Haz.
	Risk.

	D.2
	Energy hazards and contributory factors

	D.2.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.2
	Heat
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.3
	Mechanical force
	Headbox
	1
	2
	2
	Recommendation not to force doors etc when adjusting.  If damaged, user to assess level of damage / sharp edges before re-use.
	F. User manual / label
	1
	2
	2

	D.2.4
	Ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.5
	Non ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.6
	Moving parts
	Headbox
	1
	2
	2
	Recommendation that care should be taken when adjusting doors etc.
	F. User manual / label
	1
	2
	2

	D.2.7
	Unintended motion
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.8
	Suspended masses
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.9
	Patient support failure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.10
	Pressure (vessel rupture)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.11
	Acoustic pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.12
	Vibration
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.13
	Magnetic fields 

(eg. MRI)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3
	Biological hazards and contributory factors

	D.3.1
	Bio-contamination
	Headbox
	1
	2
	2
	Construction / polished surfaces – easy to clean
	E. Risk Analysis Report
	1
	2
	2

	D.3.2
	Bio-incompatibility
	Headbox
	1
	1
	1
	Perspex / acrylic sheet, rubber gaiter on neck door
	Manufacturers data
	1
	1
	1

	D.3.3
	Incorrect formulation (chemical composition)
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.4
	Toxicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.5
	Allergenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.6
	Mutagenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.7
	Oncogenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.8
	Carcinogenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.9
	Re and/or cross infection
	Headbox
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.10
	Pyrogenicity
	N/A
	1
	1
	1
	
	Manufacturers data
	1
	1
	1

	D.3.11
	Inability to maintain hygienic standards
	
	4
	2
	8
	Construction / polished surfaces – easy to clean Cleaning recommendation in user manual / label
	F. User manual / label

E. Risk Analysis Report
	1
	2
	2

	D.3.12
	Degradation
	Headbox
	1
	1
	1
	Care instructions given in the user manual
	F. User manual
	1
	1
	1

	D.4
	Environmental hazards and contributory factors

	D.4.1
	Electromagnetic fields
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.2
	Susceptibility to electromagnetic interference
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.3
	Emissions of electromagnetic interference 
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.4
	Inadequate supply of power 
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.5
	Inadequate supply of coolant
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.6
	Storage / operation outside prescribed environmental conditions
	Headbox
	1
	2
	2
	Environmental storage / operating conditions in user manual / labelling
	(F) User manual / label

(E) Risk Analysis Report
	1
	2
	2

	D.4.7
	Incompatibility with other devices with which the product is intended to be used
	O2 sensors, IV’s, tubes, leads
	1
	1
	1
	Suitably sized cutouts / holes used. 
	
	1
	1
	1

	D.4.8
	Accidental mechanical damage
	Headbox
	1
	1
	1
	Relatively robust material used. Mechanical as well as glued joints used. If damaged, user to assess level of damage / sharp edges before re-use.
	F. user manual / label
	1
	1
	1

	D.4.9
	Contamination due to waste products and/or device disposal
	Headbox
	1
	2
	2
	No special disposal required.
	Manufacturers data.

F. User manual / label

E. Risk Analysis Report
	1
	2
	2

	D.5
	Hazards resulting from incorrect output of energy and substances

	D.5.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.2
	Radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.3
	Volume
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.4
	Pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.5
	Supply of medical gases
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.6
	Supply of anaesthetic agents
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6
	Hazards related to the use of the medical device and contributory factors

	D.6.1
	Inadequate labelling
	User manual / label
	2
	1
	2
	Product easy to use

Label
	F. Label
	1
	1
	1

	D.6.2
	Inadequate operating instructions
	User manual
	2
	1
	2
	Product easy to use

User manual 
	F. User manual 
	1
	1
	1

	D.6.3
	Inadequate specification of accessories
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.4
	Inadequate specification of pre-use checks
	User manual / label
	2
	2
	4
	Product easy to use

User manual & inserts
	F. User manual / label
	2
	1
	2

	D.6.5
	Over-complicated operating instructions
	User manual
	2
	1
	2
	Product easy to use

User manual / label
	F. User manual / label
	1
	1
	1

	D.6.6
	Inadequate specification of service and maintenance
	N/A


	1
	1
	1
	No service required except external cleaning
	F. User manual / label


	1
	1
	1

	D.6.7
	Use by unskilled / untrained personnel
	Headbox
	2
	1
	2
	Product easy to use

User manual / label
	F. User manual / label

E. Risk Analysis Report
	2
	1
	2

	D.6.8
	Reasonable foreseeable misuse
	Headbox
	1
	1
	1
	Product easy to use

User manual / label
	F. User manual / label

E. Risk Analysis Report
	1
	1
	1

	D.6.9
	Insufficient warning of side effects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.10
	Inadequate warnings of hazards likely with re-use of single use devices
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.11
	Incorrect measurement and other metrological aspects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.12
	Misrepresentation of results
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.13
	Incompatibility with consumables / accessories / other devices
	Headbox
	1
	1
	1
	Suitably sized cut outs / holes used.
	
	1
	1
	1

	D.6.14
	Sharp edges or points
	Headbox
	2
	2
	4
	If damaged, user to assess level of damage / sharp edges before re-use.
	F. User manual / label

E. Risk Analysis Report
	2
	1
	2

	D.7
	Inappropriate, inadequate or overcomplicated user interface ( man/machine communication)

	D.7.1
	Mistakes & judgement errors
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.2
	Lapses and cognitive recall errors
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.3
	Slips & blunders (mental or physical)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.4
	Violation or abbreviation of instructions,  procedures etc
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.5
	Complex or confusing control system
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.6
	Ambiguous or unclear device state
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.7
	Ambiguous or unclear presentation of settings, measurement, or other information
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.8
	Misrepresentation of results
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.9
	Insufficient visibility, audibility or tactility
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.10
	Poor mapping of controls to action or of displayed information to actual state
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.7.11
	Controversial modes or mappings as compared to existing equipment
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8
	Hazards arising from function failure, maintenance and ageing and contributory factors

	D.8.1
	Erroneous data transfer
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.2
	Lack of, or inadequate specification for maintenance including post maintenance functional tests
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.4
	Inadequate maintenance
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.5
	Lack of adequate determination of end of device life
	Headbox
	1
	1
	1
	User decision based on clarity of plastics & tolerable damage to headbox.
	E. Risk Analysis Report
	1
	1
	1

	D.8.6
	Loss of electrical integrity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.7
	Loss of mechanical integrity
	Headbox
	1
	1
	1
	User decision based on clarity of plastics & tolerable damage to headbox.
	E. Risk Analysis Report
	1
	1
	1

	D.8.8
	Inadequate packaging (contamination and / or deterioration of the device)
	Headbox
	2
	2
	4
	Often exact design specified by customer -product hand packed for despatch by courier
	L. Packaging

E. Risk Analysis Report
	1
	1
	1

	D.8.9
	Re-use and/or improper re-use
	Headbox
	1
	1
	1
	User decision on suitability for next use based on clarity of plastics & tolerable damage to headbox.
	E. Risk Analysis Report
	1
	1
	1

	D.8.10
	Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
	Headbox
	1
	1
	1
	User decision based on clarity of plastics & tolerable damage to headbox.
	E. Risk Analysis Report
	1
	1
	1



T:\Simon Watmough\Projects\CE files\Headboxes\E. Risk Management BS14971 Annex D_Headboxes_SW.doc
    Page 7
02/04/03
Severity of hazard : 1 = Insignificant, 2 =  Tolerable, 3  = Critical, 4  =  Intolerable.  Probability of event : 1 = Improbable, 2 = Occasional, 3 = Likely, 4 = Highly likely.  

Risk  calculated as severity of hazard x probability of event, 1 – 16.  Further explanation of risk management policy – see Risk Management Policy & Definitions.


