[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




[image: image2.png]



                                                         Headboxes

Risk Analysis

This product range was designed by hospitals in the early 1960’s and has been continually improved in line with medical practices

Headboxes are small enclosures usually used to surround a neonate’s head to allow a greater concentration of Oxygen to be achieved. They are manufactured from transparent 4mm or 5mm Acrylic (Perspex) for biocompatibility and lack of toxic risk to newborns. and are shaped with a neck opening.

Several “standard sizes” are available to cover the range of patients Neonate; infant, Paediatric.

Specials are manufactured to the users specification but all follow the basic description. The main requirements for change being size and position of the access holes.

They all have sliding neck doors to enable a better seal around the patients neck.

The Neonatal neck door is slanted backward into the Headbox to allow room for the

Neonates chin. This avoids abrasion with the shoulders as the infant moves.

Half holes are drilled around the base to give access for IV tubes and leads etc. this prevents instability

All headboxes have at least a 22mm hole for Oxygen.

Some versions have rotating side windows,

C02 can build up if there is no flow or a very low flow

All edges of the device are polished in the manufacturing process to prevent abrasion to the patient

The shape is a squared hemi-sphere without sharp corners to make cleaning and disinfection easy to achieve, to minimise cross-infection. The flat panels are designed to maximise visibility of the patient and reduce distortion

WARNINGS

To avoid carbon dioxide accumulation a warning to use gas flows of at least 6 litres/mm is a label fixed to the Headbox

b)Instruction Sheet

It is recommended that when used with Neonates the headbox should be pre-warmed to prevent cooling of the infant from radiant losses to the acrylic.

It is also recommended that gas be warned and humidified before introduction into the headbox is to prevent cooling of the infant, and increased fluid losses from the skin.

W:\COMPANY\CEMARK\PLastics WP\Headboxes\(E) Risk report Headboxes.doc
Page 1
December 18, 2002

[image: image1.jpg][image: image2.png]