[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




[image: image2.png]




MASTER– CE FILE

II. IMPLEMENTING RULES.

2.
Implementing rules.
2.1.
Application of the classification rules shall be governed by the intended purpose of the devices.

2.2.
If the device is intended to be used in combination with another device, the classification rules shall apply separately to each of the devices.  Accessories are classified in their own right separately from the device with which they are used.

2.3.
Software, which drives a device or influences the use of a device, falls automatically in the same class.

2.4.
If the device is not intended to be used solely or principally in a specific part of the body, it must be considered and classified on the basis of the most critical specified use.

2.5.
If several rules apply to the same device, based on the performance specified for the device by the manufacturer, the strictest rules resulting in the higher classification shall apply.

III.  CLASSIFICATION.

1.
Non-invasive devices
1.1 All Non-invasive devices are in Class I, unless one of the rules set out hereinafter applies:

3. Additional Rules Applicable to Devices

3.2 Rule 10.


Active devices intended for diagnosis are in Class II(a):

· if they area intended to supply energy which will be absorbed y the human body, except for devices used to illuminate the patient’s body, in the visible spectrum.

· If they are intended to image in vivo distribution of radio pharmaceuticals,
[image: image1.jpg]
Active devices intended to emit ionising radiation and intended for diagnostic and therapeutic interventional radiology including devices which control or monitor such devices, or which directly influence their performance, are in Class II(b).







If they are intended to allow direct diagnosis or monitoring of vital physiological processes, unless they are specifically intended for monitoring of vital physiological parameters, where the nature of variations is such that it could result in immediate danger to the patient, for instance variations in cardiac performance, respiration, activity of CNS in which case they are in Class II(b).











Classification Implementing the rules.doc                         27/09/2006    Page 1 of 1

[image: image2.png]