V1.1    01/06/11

Production Procedure

Oxygen Analyzer assembly 7900000 

Potentiometer assembly

Cut red wire (7930009) and black wire (7930010) to 6.5cm, strip 0.5cm from both ends of each and tin the wires.

Solder red (LHS) and black (RHS) wires to the sensor connector. Fit 1.0cm strips of heat shrink tubing (0032310) to both individual cables, slide over the terminals and shrink the heat shrink tubing. Then fit 1.5cm strip of heat shrink (0032312) over both the cables, shrink the heat shrink tubing.

Cut red wire (7930009) to 11.0cm and strip back 0.5cm from both ends, tin the wires.

Cut black wire (7930010)  to 10.00cm and strip back 0.5cm from both ends, tin the wires.

Solder red – connector wire to the potentiometer LHS terminal.

Solder the red loose cable to the upper centre potentiometer terminal.

Solder the black – connector wire and the black loose wire to the potentiometer RHS terminal.

Use 1.00cm strips (0032312 LHS & centre and 0032310 to RHS) of heat shrink to cover the potentiometer terminals, shrink the heat shrink.
LCD preparation

Short link J1 using solder.

Using tinned copper wire short together:

a) Two terminals: VDD (2nd from LHS) and DP2 (4th from LHS). Ensure that the link doesn’t touch DP1 solder pad.

b) Four terminals: COM, RFL, INLO, ROL.
c) Two terminals: RFH, ROH. 

NOTE: Fit the shorting links to the display side (underside) of the PCB.
Case preparation

Stick black sealing strip (7930015) to the grey strip (7930014) and cut to 3.2cm lengths. Fit sealing strips to the battery compartment.

Fit switch to the case (0 to the facing bottom of the case) – use primer and cyanoacrylate to secure. Allow 1/2hr to dry.

Fit oxygen sensor adapter to the front of the case – use primer and cyanoacrylate. Allow 1/2hr to dry. NOTE: use an old oxygen sensor to aid alignment and to secure whilst drying.

Battery cable assembly

Securely affix cable tie at 4.0cm from the connector.

Cut red cable at 8.5cm and strip wire back by 0.3cm, tin the wire.

Fit 1.0cm strip of heat shrink tubing (0032312) onto the red wire and solder the red wire to the switch  (lower terminal). Slide the heat shrink over the terminal and shrink.

Final assembly

Fit potentiometer assembly.

Solder wires to the LCD:

Black battery lead to VSS
Red switch wire to VDD
Black potentiometer assembly wire to INLO
Red potentiometer wire to INHI
Using test sensor ensure that the unit is working correctly. Fit and secure LCD using bead of silicone adhesive (7950000). Allow assembly to dry for >12hrs.

Fit two cable saddles (7930013) to the case, one aligned with DP2 terminal and the other aligned with the RFH terminal. Secure the cable assemblies in the cable saddles.

Fit cable ties (7930012) to cable assemblies:

a) Two to the potentiometer cable.

b) Two to the battery cable.

c) One between the two cable saddles

Affix serial number label (7930019) to the case rear panel.

Assemble case and carry out functional test.

Clean case with silicone polish and place in plastic bag. Affix tracking label to plastic bag (XX).
