	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Conversion  AD300/MD300 to Mains PSU

	VM3/COP/40.55

	Date: 23 jan- 2009
	Revision Date: 17-Jun-11
	Issue 1

	Only to be attempted by trained personel



Parts Required

	Quantity
	Description
	Part No.
	

	1
	AD or MD300
	
	

	1
	Mains PSU Mixcheck
	
	

	1
	Wire black
	
	

	1
	21.mm Power connector
	9071160
	

	1
	21.mm Rt Angle plug
	
	

	1
	Silicone
	9711025
	

	2
	 Heat shrink 1cm long
	
	


[image: image1.jpg]


[image: image2.jpg]


Procedure

1.  Remove the battery cover

2. Drill a 2mm in the centre of the  middle band of battery locators

3. Using increasing diameter drill bits increase the hole size to 10mm

4. Using the tape bit increase the hole size by

hand until the socket fits the hole.

5. Fasten the socket into the hole so that the

input is near the bottom of the instrument

6. Add glue to the thread and the join between the nut and the case to prevent the socket turning whilst in us.

7. Bend the terminals at right angles to reduce the height if the socket in the enclosure.

8. Remove the back cover holding screws (5 in the battery compartment and one at the top of the instrument)

9. [image: image3.jpg]


[image: image4.jpg]


[image: image5.jpg]


Drill two 2 mm holes one each side if the bottom battery terminals

10. [image: image6.jpg]


Feed a red wire through the right hole  and a Black or Blue wire through the left hole 

11. Leave wires approximately 10 cm. long  Make small loops in the wire and solder

12. Place the loops between the enclosure and the battery connector Blue to left -ve and Red to right +ve

13. Add the insulator and use screws to hold whilst the case is reassembled

14. Ensure the gasket is in place 

15. Check continuity Red wire to +ve terminal and Black/Blue to -ve terminal

16. Use a standard Mixcheck PSU 

17. The outer is -ve and the inner is +ve

18. Cut off the straight connector and add a right angle connector


19. [image: image7.jpg]


[image: image8.jpg]


[image: image9.jpg]


Ensure the cable is soldered on to the connector +ve to centre -ve to outer

20. Use the clamp to clamp cable

21. Pot the connector in Silcone

22. Wire the socket Red to centre Black/Blue to outer and heat-shrink

23. Re-assemble the battery cover

24. Ensure the gasket is in place

25. Check Instrument functions OK

VM3COP40.55.doc       17/06/2011      Page 2 of 2

